## Methods by Batch
|  | 1999 Eggs 1-30 | 2000 Eggs 31-67 | 2001 Eggs 68-87 | 2002 Eggs 123-139; 180-188; 192-199 | 2003 Eggs 297-336; 2004 Eggs 464-513; 2005 Eggs 676-683; 748-871 | 2005 Eggs 576, 578, 579, 611, 614, 616, 617, 619, 622, 641, 642, 644, 655, 658, 661, 664, 666, 669, 670, 673, 674 | 2008 Eggs 1234-1416 | 2009 Eggs 1419 - 1486 | 2010 Eggs 1487 - 1619 | Temporal Eggs 2002 - 114-120; 189-191; 2003 - 275-284; 2004 - 406-418; 455-462; 472-476; 2005 - 845-853; 2006 - 985-1018; 1069-1100; 1127-1136
Prep | Homogenization | Cryo | Cryo | Cryo | Cryo | Cryo | Fresh | Fresh | Fresh | Fresh | Fresh
| Storage | Teflon jars | Teflon jars | Teflon jars | Teflon jars | Teflon jars | Teflon jars and polypropylene cryovials | Teflon jars and polypropylene cryovials | Teflon jars and polypropylene cryovials | Teflon jars and polypropylene cryovials | Teflon jars and polypropylene cryovials
Organic analysis (BDEs, PCBs, organochlorine pesticides) | Calibration | SRMs 2261,2262,2274,2275 PCB solution 3 | SRMs 2261,2262,2274,2275 PCB solution 3 | SRMs 2261 (Chlorinated Pesticides in Hexane), 2262 (Chlorinated Biphenyl Congeners in Isooctane), 2274 (PCB Congener Solution-II in Isooctane), 2275 (Chlorinated Pesticide Solution-II in Isooctane), PCB Congener Solution III, and PCB Congener Solution IV | SRMs 2261 (Chlorinated Pesticides in Hexane), 2262 (Chlorinated Biphenyl Congeners in Isooctane), 2274 (PCB Congener Solution-II in Isooctane), 2275 (Chlorinated Pesticide Solution-II in Isooctane), PCB Congener Solution III, and PCB Congener Solution IV | SRMs 2257 (Brominated Diphenyl Ethers in Isooctane), 2258 (BDE 209 in Isooctane), and 2259 (Chlorinated Biphenyl Congeners in Isooctane), 2261 (Chlorinated Pesticides in Hexane), and 2275 (Chlorinated Pesticide Solution-II in Isooctane), and solutions of HBCDs, octachlorostyrene, and pentachlorobenzene | SRM 2261 (Chlorinated Pesticides in Hexane), SRM 2262 (Chlorinated Biphenyl Congeners in Isooctane), SRM 2274 (PCB Congener Solution-II in Isooctane), SRM 2275 (Chlorinated Pesticide Solution-II in Isooctane), PCB Congener Solution III, PCB Congener Solution IV, PBDE Congener Solution containing 26 congeners, BDE 209 solution, and α-, β-, and γ-HBCDD solution | SRMs 2257 (Brominated Diphenyl Ethers in Isooctane), 2259 (Chlorinated Biphenyl Congeners in Isooctane), 2261 (Chlorinated Pesticides in Hexane), and 2275 (Chlorinated Pesticide Solution-II in Isooctane), and solutions of HBCDs, octachlorostyrene, and pentachlorobenzene | SRMs 2257 Brominated Diphenyl Ethers in Isooctane, 2259 Chlorinated Biphenyl Congeners in Isooctane, 2261 Chlorinated Pesticides in Hexane, and 2275 Chlorinated Pesticide Solution-II in Isooctane, and solutions of HBCDs, octachlorostyrene, and pentachlorobenzene | SRMs 2257 Brominated Diphenyl Ethers in Isooctane, 2259 Chlorinated Biphenyl Congeners in Isooctane, 2261 Chlorinated Pesticides in Hexane, and 2275 Chlorinated Pesticide Solution-II in Isooctane, and solutions of octachlorostyrene and pentachlorobenzene | SRMs 2257 Brominated Diphenyl Ethers in Isooctane, 2259 Chlorinated Biphenyl Congeners in Isooctane, 2261 Chlorinated Pesticides in Hexane, and 2275 Chlorinated Pesticide Solution-II in Isooctane, and solutions of octachlorostyrene and pentachlorobenzene
| Internal Standard | 4,4’-DDTd8, 4,4’-DDE-d8, 4,4’-DDD-d8, endosulfan I-d4, PCB 103 and PCB 198, PCB 77, PCB 126 and PCB 169 | 4,4’-DDTd8, 4,4’-DDE-d8, 4,4’-DDD-d8, PCB 103 and PCB 198, PCB 77, PCB 126 and PCB 169 | 13C labeled PCB congeners 28, 52, 118, 153, 180, 194, and 206; deuterated 4,4’-DDD (d8), 4,4’-DDT (d8), 4,4-DDE (d8), and endosulfan-I (d4); and 13C labeled trans-chlordane | 13C labeled PCB congeners 28, 52, 118, 153, 180, 194, and 206; deuterated 4,4’-DDD (d8), 4,4’-DDT (d8), 4,4-DDE (d8), and endosulfan-I (d4); and 13C labeled trans-chlordane | 13C labeled PCB congeners 28, 52, 118, 153, 180, 194, and 206; 13C labeled 4,4-DDE, 4,4’-DDT, HCB, dieldrin, oxychlordane, trans-chlordane, trans-nonachlor, deuterated 4,4’-DDD (d8); F labeled BDE congeners 47 and 160; and BDE 104 | 13C labeled PCB congeners 28, 52, 77, 118, 126, 153, 169, 180, 194, and 206, BDE congeners 28, 47, 99, 100, 153, 154 and 209 and 4,4-DDE, 4,4’-DDT, HCB, dieldrin, oxychlordane, trans-chlordane, trans-nonachlor, α-HBCDD, γ-HBCDD; and 4,4’-DDD-d8, and F-PBDE 208 | 13C labeled PCB congeners 28, 52, 118, 153, 180, 194, and 206; 13C labeled 4,4-DDE, 4,4’-DDT, HCB, oxychlordane, trans-chlordane, trans-nonachlor, deuterated 4,4’-DDD (d8); F labeled BDE congeners 47 and 160; and BDE 104 | 13C labeled PCB congeners 28, 52, 118, 153, 180, 194, and 206; 13C labeled 4,4-DDE, 4,4’-DDD, 4,4’-DDT, HCB, oxychlordane, trans-chlordane, trans-nonachlor; F labeled BDE congeners 47 and 160; and BDE 104 | 13C labeled PCB congeners 28, 52, 77, 118, 126, 153, 169, 180, 194, and 206; 13C labeled 4,4-DDE, 4,4’-DDD, 4,4’-DDT, HCB, oxychlordane, trans-chlordane, trans-nonachlor, methyl triclosan, α-HBCDD, β-HBCDD, γ-HBCDD and BDE 209; F labeled BDE congeners 47, 160, and 208; and BDE 104 | 13C labeled PCB congeners 28, 52, 77, 118, 126, 153, 169, 180, 194, and 206; 13C labeled 4,4-DDE, 4,4’-DDD, 4,4’-DDT, HCB, oxychlordane, trans-chlordane, trans-nonachlor, methyl triclosan, α-HBCDD, β-HBCDD, γ-HBCDD and BDE 209; F labeled BDE congeners 47, 160, and 208; and BDE 104
| Extraction | PFE | PFE | PFE | PFE | PFE | PFE | PFE | PFE | PFE | PFE
| Sorbent | Sodium sulfate | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth | Diatomaceous Earth
| Clean-up | SEC (PLGel 600 mm x 25 mm -10 µm particle size with 100 Å diameter pores), semi-prep aminopropylsilane fractionation | SEC (PLGel 600 mm x 25 mm -10 µm particle size with 100 Å diameter pores), semi-prep aminopropylsilane fractionation | SEC (PLGel 600 mm x 25 mm -10 µm particle size with 100 Å diameter pores), SPE with alumina | SEC (PLGel 600 mm x 25 mm -10 µm particle size with 100 Å diameter pores), SPE with alumina | SEC (Phenogel 50 mm x 7.80 mm i.d; 10 µm particle size guard column coupled to a 600 mm x 21.2 mm Phenogel column in series with a Phenogel 300 mm x 21.2 mm (10 µm particle size with 100 Å diameter pores); SPE with alumina | SEC (PLGel 110 mm x 25 mm i.d; 10 µm particle size guard column coupled to a 600 mm x 25 mm (10 µm particle size with 100 Å diameter pores) PLGel column); SPE with alumina | SEC (PLGel 110 mm x 25 mm i.d; 10 µm particle size guard column coupled to a 600 mm x 25 mm (10 µm particle size with 100 Å diameter pores) PLGel column); SPE with alumina | SEC (PLGel 110 mm x 25 mm i.d; 10 µm particle size guard column coupled to a 600 mm x 25 mm (10 µm particle size with 100 Å diameter pores) PLGel column); SPE with alumina | SEC (PLGel 110 mm x 25 mm i.d; 10 µm particle size guard column coupled to a 600 mm x 25 mm (10 µm particle size with 100 Å diameter pores) PLGel column); SPE with alumina | SEC (PLGel 110 mm x 25 mm i.d; 10 µm particle size guard column coupled to a 600 mm x 25 mm (10 µm particle size with 100 Å diameter pores) PLGel column); SPE with alumina
| Analysis | GC-ECD with 60 m x 0.25 mmx0.25 µm DB-5 & DB-XLB columns with 5 m x 0.25 mm guard columns | GC-ECD with 60 m x 0.25 mmx0.25 µm DB-5 & DB-XLB columns with 5 m x 0.25 mm guard columns | GC/MS with 60 m x 0.25 mmx0.25 µm DB-5MS with 5 m x 0.25 mm guard column | GC/MS with 30 m x 0.18 mmx 0.18 µm DB-XLB with 5 m x 0.25 mm guard column | EI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-5MS and NCI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-XLB with 5 m x 0.25 mm guard columns | EI-GC/MS with 60 m x 0.25 mmx 0.25 µm DB-5MS and NCI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-5MS with 5 m x 0.25 mm guard columns | EI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-5MS and NCI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-XLB with 5 m x 0.25 mm guard columns | EI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-5MS and NCI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-XLB with 5 m x 0.25 mm guard columns | EI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-5MS and NCI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-XLB with 5 m x 0.25 mm guard columns | EI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-5MS and NCI-GC/MS with 30 m x 0.18 mmx 0.18 µm DB-XLB with 5 m x 0.25 mm guard columns
Mercury | Extraction | Nitric acid closed vessel microwave | Nitric acid closed vessel microwave | Nitric acid closed vessel microwave | Nitric acid in open focus microwave | Nitric acid in open focus microwave | Nitric acid in open focus microwave | Nitric acid and hydrogen peroxide closed vessel microwave | Nitric acid and hydrogen peroxide in open focus microwave | - | -
| Analysis | ID-CV-ICPMS | ID-CV-ICPMS | ID-CV-ICPMS | ID-CV-ICPMS | ID-CV-ICPMS | ID-CV-ICPMS | ID-CV-ICPMS | ID-CV-ICPMS | Direct Mercury Analysis | Direct Mercury Analysis

## Organics
|  | 1999 - 2000 eggs: 1-67 | [13 empty] | 2002&2003 eggs 123-139;180-198; PeFa |  |  |  | 2003-2005 eggs - 297 - 336; 464-513; 676-683; 748-871 | [8 empty] | 2005 gull eggs 576, 578, 579, 611, 614, 616, 617, 619, 622, 641, 642, 644, 655, 658, 661, 664, 666, 669, 670, 673, 674 |  |  |  |  | 2008 eggs 1234-1416 |  |  |  |  | 2009 eggs 1419 - 1486 | [18 empty] | 2010 STAMP eggs 1487 - 1619 | [21 empty] | Temporal - 2002 eggs 114-120; 189-191; 2003 - 275-284; 2004 - 406-418; 455-462; 472-476; 2005 - 845-853; 2006 - 985-1018; 1069-1100; 1127-1136
|  |  | SRM 1946 | [6 empty] | Certified Values | Reference Values | 2001 eggs 68-87 | [8 empty] | SRM 1946 Certified | SRM 1946 |  |  | Murre Egg CM | Murre Egg CM |  |  |  | SRM 1946 Certified | SRM | Murre Egg CM | Murre | [8 empty] | Murre Egg Control Material | [15 empty] | ng/g wet wt after LOD: | CM consensus | EggCM-1.D | EggCM-2.D | EggCM-3.D | SRM 1946 Cert/Ref | 1946-1.D | 1946-2.D | 1946-3.D | [15 empty] | Murre Egg Control Material
|  |  | 1,3,12,20,24 | 2,4,5,10,11,13,16,21-23 | 6-9,14-15,17-19,25-30 | 31-41 | 42-67 | [14 empty] | 297-513 |  | 676-871 |  | 297-513 |  | 676-871
ng/g wet wt: |  | Compound | Batch 1 | Batch 2 | Batch 3 | Batch 4 | Batch 5 | Mean | SD | Mean + SD | Mean + SD | 1946 | MURRE | Compound | SRM 1946 Certified or Reference Values | 1946 | Murre Egg CM Reference Values | Batch 1 | Batch 2 | Compound | or Reference Values | Batch 1 | Batch 2 | Batch 3 | Reference Values | Batch 1 | Batch 2 | Batch 3 | Compound | or Reference Values | 1946 | Reference (CHS) Values | Egg CM | Compound |  | Consensus Values |  | Murre Egg CM | Compound | Max LOD or LOQ |  | Consensus Values |  | Batch 1-1 | Batch 1-2 | Batch 1-3 |  | Mean | SD | RSD | [29 empty] | Compound | Max LOD or LOQ | Consensus Values | Batch 1-1 | Batch 1-2 | Batch 1-3 | Batch 2-1 | Batch 2-2 | Batch 2-3
|  | % lipid | 8.63 | 9.44 | 11.8 | 7.43 | 8.32 | 9.124 | 1.6599 | 10.17 + 0.48 |  |  |  | percent lipid | 10.17 ± 0.48 | 9.56 | 11.8 ± 1.1 | 10.9 | 10.3 | [9 empty] | percent lipid | 10.17 ± 0.48 | 9.5604 | 11.8 ± 1.1 | 10.9373 | [5 empty] | % lipid | - | - |  | 11.8 ± 1.1 |  | 8.75 | 8.62 | 8.95 |  | 8.77 | 0.17
PCB 18 |  | PCB 18 | <0.100 | <0.100 | <0.100 | 0.27 | 0.33 | 0.15 | 0.14 |  | 0.84 + 0.11 | 0.34242 | 0.05 | PCB 18 | 0.84 + 0.11 | 0.716 | <0.100 | <0.100 | <0.100 | PCB 18 | 0.84 + 0.11 | 0.449 |  | 0.422 | <0.100 | <LOQ |  | <LOQ | PCB 18 | 0.84 + 0.11 | 0.466 | <0.100 | <0.100 | PCB 18 |  | <0.100 |  | <LOD | PCB 18 | 0.0065 |  | <0.100 |  | <LOD | <LOD | <LOD |  | <LOD | [9 empty] | PCB 18 | <0.100 | 0.022 | 0.023 | 0.021 | 0.84 + 0.11 | 0.323 | 0.321 | 0.326 | [13 empty] | PCB 18 | 0.018 | <0.100 | <LOD | <LOD | <LOD
PCB28+31 |  | PCB 28 | 2.14 | 1.96 | 1.75 | 2.04 | 2.09 | 1.996 | 0.15274 |  | 2.00 + 0.24 | 3.3875 | 2.714 | PCB28+31 | 2.00 + 0.24; 1.46 + 0.20 | 3.35 | 2.49 ± 0.51 | 2.58 | 2.53 | PCB 28+31 | 2.00 + 0.24; 1.46 + 0.20 | 2.8 |  | 2.76 | 2.49 ± 0.51 | 2.22 |  | 2.15 | PCB28+31 | 2.00 + 0.24; 1.46 + 0.20 | 3.35 | 2.49 ± 0.51 | 2.74 | PCB 28+31 |  | 2.49 ± 0.51 |  | 2.95 | PCB28+31 | 0.0204 |  | 2.49 ± 0.51 |  | 2.03 | 2.01 | 2.03 |  | 2.02 | 0.011 | 0.53 | [7 empty] | PCB28+31 | 2.49 ± 0.51 | 2.159 | 2.149 | 2.176 | 2.00 + 0.24; 1.46 + 0.20 | 2.79 | 2.833 | 2.975 | [13 empty] | PCB28+31 | 0.062 | 2.49 ± 0.51 | 2.153 | 2.132 | 2.237
PCB 29 | [19 empty] | PCB 29 |  | <LOQ |  | <LOQ | <0.100 | <LOQ |  | <LOQ | [5 empty] | PCB 29 |  | <0.100 |  | <LOQ | PCB 29 | 0.00117 |  | <0.100 |  | <LOD | <LOD | <LOD |  | <LOD | [9 empty] | PCB 29 | <0.100 | <LOQ | <LOQ | <LOQ |  | #VALUE! | #VALUE! | #VALUE!
PCB 44 |  | PCB 44 | 5.21 | 4.68 | 3.94 | 4.97 | 5.52 | 4.864 | 0.60186 | 4.66 + 0.86 |  | 3.7495 | 0.05 | PCB 44 | 4.66 + 0.86 | 4.68 | <0.100 | <0.100 | <0.100 | PCB 44 | 4.66 + 0.86 | 3.41 |  | 3.2 | <0.100 | <LOD |  | <LOD | PCB 44 | 4.66 + 0.86 | 3.95 | <0.100 | <0.100 | PCB 44 |  | <0.100 |  | 0.0577 | PCB 44 | 0.0605 |  | <0.100 |  | 0.0749 | 0.0788 | 0.0724 |  | 0.0754 | 0.0032 | 4.3 | [7 empty] | PCB 44 | <0.100 | 0.062 | 0.066 | 0.067 | 4.66 + 0.86 | 3.069 | 3.024 | 3.118 | [13 empty] | PCB 44 | 0.026 | <0.100 | 0.015 | 0.015 | 0.017
PCB 45 | [19 empty] | PCB 45 |  | 2.45 |  | 2.53 | <0.100 | 0.0482 |  | <LOQ | [5 empty] | PCB 45 |  | <0.100 |  | 0.0095 | PCB 45 | 0.00905 |  | <0.100 |  | 0.0328 | 0.0282 | 0.0307 |  | 0.0306 | 0.0023 | 7.5 | [7 empty] | PCB 45 | <0.100 | <LOD | <LOD | <LOD |  | 0.079 | 0.079 | 0.073 | [13 empty] | PCB 45 | 0.031 | <0.100 | 0.076 | 0.073 | 0.085
PCB 49 | [11 empty] | 3.2331 | 0.05 | PCB 49 | 3.8 + 0.39 | 3.56 | <0.15 | <0.100 | <0.100 | PCB 49 | 3.8 + 0.39 | 2.98 |  | 2.96 | <0.15 | 0.0646 |  | 0.0516 | PCB 49 | 3.8 + 0.39 | 3.57 | <0.15 | 0.423 | PCB 49 |  | <0.100 |  | 0.106 | PCB 49 | 0.0341 |  | <0.100 |  | 0.0875 | 0.0917 | 0.0912 |  | 0.0902 | 0.0023 | 2.5 | [7 empty] | PCB 49 | <0.100 | 0.099 | 0.096 | 0.092 | 3.8 + 0.39 | 2.631 | 2.632 | 2.682 | [13 empty] | PCB 49 | 0.031 | <0.100 | 0.054 | 0.056 | 0.059
PCB 52 |  | PCB 52 | 7.59 | 8.47 | 7.2 | 8.07 | 8.85 | 8.036 | 0.66142 | 8.1 + 1.0 |  | 8.2653 | 0.14497 | PCB 52 | 8.1 + 1.0 | 7.56 | 0.234 ± 0.095 | 0.201 | 0.128 | PCB 52 | 8.1 + 1.0 | 6.78 |  | 6.68 | 0.234 ± 0.095 | 0.134 |  | 0.127 | PCB 52 | 8.1 + 1.0 | 8.50 | 0.234 ± 0.095 | 0.290 | PCB 52 |  | 0.234 ± 0.095 |  | 0.161 | PCB 52 | 0.1 |  | 0.234 ± 0.095 |  | 0.192 | 0.194 | 0.19 |  | 0.192 | 0.002 | 1 | [7 empty] | PCB 52 | 0.234 ± 0.095 | 0.157 | 0.146 | 0.153 | 8.1 + 1.0 | 6.318 | 6.367 | 6.234 | [13 empty] | PCB 52 | 0.046 | 0.234 ± 0.095 | 0.148 | 0.158 | 0.152
PCB 56 | [11 empty] | 2.3295 | 0.52111 | PCB 56 | 5.77 + 0.93 | 2.77 | 0.746 ± 0.059 | 0.767 | 0.884 | PCB 56 | 5.77 + 0.93 | 3.07 |  | 3.7 | 0.746 ± 0.059 | 0.545 |  | 0.535 | PCB 56 | 5.77 + 0.93 | 2.04 | 0.746 ± 0.059 | 0.602 | PCB 56 |  | 0.746 ± 0.059 |  | 0.707 | PCB 56 | 0.145 |  | 0.746 ± 0.059 |  | 0.572 | 0.572 | 0.612 |  | 0.585 | 0.023 | 3.9 | [7 empty] | PCB 56 | 0.746 ± 0.059 | 0.728 | 0.731 | 0.772 | 5.77 + 0.93 | 5.504 | 5.502 | 5.445 | [13 empty] | PCB 56 | 0.084 | 0.746 ± 0.059 | 0.727 | 0.728 | 0.704
PCB 63 | [11 empty] | 1.1491 | 0.45037 | PCB 63 | 1.28 + 0.19 | 1.17 | <0.6; 0.500 ± 0.043 | 0.541 | 0.554 | PCB 63 | 1.28 + 0.19 | 1.19 |  | 1.14 | <0.6; 0.500 ± 0.043 | <LOQ |  | <LOQ | PCB 63 | 1.28 + 0.19 | 1.22 | <0.6; 0.500 ± 0.043 | 0.697 | PCB 63 |  | 0.500 ± 0.043 |  | 0.19 | PCB 63 | 0.0324 |  | 0.500 ± 0.043 |  | 0.169 | 0.17 | 0.171 |  | 0.17 | 0.00098 | 0.58 | [7 empty] | PCB 63 | 0.500 ± 0.043 | 0.447 | 0.49 | 0.466 | 1.28 + 0.19 | 1.27 | 1.356 | 1.348 | [13 empty] | PCB 63 | 0.183 | 0.500 ± 0.043 | 0.255 | 0.222 | 0.227
PCB 66 |  | PCB 66 | 11.1 | 10.4 | 8.74 | 9.8 | 10.2 | 10.048 | 0.86978 | 10.8 + 1.9 |  | 9.5767 | 2.2084 | PCB 66 | 10.8 + 1.9 | 10.5 | 2.41 ± 0.12 | 1.91 | 2.22 | PCB 66 | 10.8 + 1.9 | 7.94 |  | 7.98 | 2.41 ± 0.12 | 1.88 |  | 1.75 | PCB 66 | 10.8 + 1.9 | 9.58 | 2.41 ± 0.12 | 2.34 | PCB 66 |  | 2.41 ± 0.12 |  | 2.32 | PCB 66 | 0.146 |  | 2.41 ± 0.12 |  | 1.81 | 1.84 | 1.94 |  | 1.87 | 0.069 | 3.7 | [7 empty] | PCB 66 | 2.41 ± 0.12 | 2.358 | 2.371 | 2.355 | 10.8 + 1.9 | 8.977 | 8.991 | 9.039 | [13 empty] | PCB 66 | 0.058 | 2.41 ± 0.12 | 2.351 | 2.313 | 2.317
PCB 70 | [11 empty] | 10.4437 | 0.15711 | PCB 70 | 14.9 + 0.6 | 13.0 | 0.205 ± 0.018 | 0.169 | 0.170 | PCB 70 | 14.9 + 0.6 | 10.5 |  | 9.94 | 0.205 ± 0.018 | 0.136 |  | 0.137 | PCB 70 | 14.9 + 0.6 | 11.2 | 0.205 ± 0.018 | 0.189 | PCB 70 |  | 0.205 ± 0.018 |  | 0.121 | PCB 70 | 0.0491 |  | 0.205 ± 0.018 |  | 0.118 | 0.126 | 0.127 |  | 0.124 | 0.005 | 4.1 | [7 empty] | PCB 70 | 0.205 ± 0.018 | 0.187 | 0.189 | 0.19 | 14.9 + 0.6 | 14.672 | 14.407 | 14.494 | [13 empty] | PCB 70 | 0.085 | 0.205 ± 0.018 | 0.203 | 0.193 | 0.206
PCB 74 | [11 empty] | 4.8479 | 1.7847 | PCB 74 | 4.83 + 0.51 | 4.81 | 1.92 ± 0.061 | 1.78 | 1.96 | PCB 74 | 4.83 + 0.51 | 4.09 |  | 4.24 | 1.92 ± 0.061 | 1.48 |  | 1.41 | PCB 74 | 4.83 + 0.51 | 4.53 | 1.92 ± 0.061 | 1.73 | PCB 74 |  | 1.92 ± 0.061 |  | 1.97 | PCB 74 | 0.0969 |  | 1.92 ± 0.061 |  | 1.41 | 1.42 | 1.49 |  | 1.44 | 0.041 | 2.9 | [7 empty] | PCB 74 | 1.92 ± 0.061 | 1.902 | 1.977 | 1.936 | 4.83 + 0.51 | 4.997 | 4.931 | 5.05 | [13 empty] | PCB 74 | 0.036 | 1.92 ± 0.061 | 1.883 | 1.858 | 1.862
[20 empty] | PCB 79 |  | 0.0564 |  | 0.0354 | 0.165 ± 0.026 | <LOQ |  | <LOQ | [5 empty] | PCB 79 |  | 0.165 ± 0.026 |  | 0.0221 | PCB 79 | 0.0273 |  | 0.165 ± 0.026 |  | 0.0302 | 0.0314 | 0.0266 |  | 0.0294 | 0.0025 | 8.4 | [7 empty] | PCB 79 | 0.165 ± 0.026 | <LOD | <LOD | <LOD |  | 2.996 | 2.98 | 2.803
[20 empty] | PCB 82 |  | 0.115 |  | 0.122 | <0.100 | <LOD |  | 0.0382 | [5 empty] | PCB 82 |  | <0.100 |  | 0.0284 | PCB 82 | 0.0306 |  | <0.100 |  | 0.0383 | 0.0348 | 0.0359 |  | 0.0363 | 0.0018 | 5 | [7 empty] | PCB 82 | <0.100 | 0.02 | 0.021 | 0.022 |  | 0.068 | 0.066 | 0.066 | [13 empty] | PCB 82 | 0.023 | <0.100 | <LOQ | <LOQ | <LOQ
PCB 87 | [11 empty] | 10.2731 | 0.074827 | PCB 87 | 9.4 + 1.4 | 10.9 | <0.15; 0.106 ± 0.0051 | 7.40 | 5.04 | PCB 87 | 9.4 + 1.4 | 9.82 |  | 10.2 | <0.15; 0.106 ± 0.0051 | 0.101 |  | 0.0937 | PCB 87 | 9.4 + 1.4 | 14.1 | <0.15; 0.106 ± 0.0051 | 0.203 | PCB 87 |  | 0.106 ± 0.0051 |  | 0.088 | PCB 87 | 0.0121 |  | 0.106 ± 0.0051 |  | 0.107 | 0.104 | 0.102 |  | 0.105 | 0.0024 | 2.3 | [7 empty] | PCB 87 | 0.106 ± 0.0051 | 0.109 | 0.104 | 0.102 | 9.4 + 1.4 | 10.198 | 10.166 | 10.507 | [13 empty] | PCB 87 | 0.033 | 0.106 ± 0.0051 | 0.103 | 0.109 | 0.103
[20 empty] | PCB 92 |  | 8.18 |  | 7.55 | <0.100 | <LOQ |  | 0.0155 | [5 empty] | PCB 92 |  | <0.100 |  | 0.0772 | PCB 92 | 0.0886 |  | <0.100 |  | 0.095 | 0.0953 | 0.0922 |  | 0.0942 | 0.0017 | 1.8 | [7 empty] | PCB 92 | <0.100 | 0.094 | 0.086 | 0.09 |  | 5.917 | 6.08 | 5.865 | [13 empty] | PCB 92 | 0.081 | <0.100 | 0.053 | 0.046 | 0.053
PCB 95 | [11 empty] | 10.8474 | 0.091066 | PCB 95 | 11.4 + 1.3 | 15.6 | <0.3 | <0.100 | <0.100 | PCB 95+121 | 11.4 + 1.3 | 12.5 |  | 12.6 | <0.3 | 0.0652 |  | 0.0643 | PCB 95 | 11.4 + 1.3 | 14.1 | <0.3 | 0.187 | PCB 95+121 |  | <0.100 |  | <LOD | PCB 95+121 | 0.0203 |  | <0.100 |  | 0.0628 | 0.0571 | 0.0625 |  | 0.0608 | 0.0032 | 5.3 | [7 empty] | PCB 95+121 | <0.100 | 0.071 | 0.074 | 0.073 | 11.4 + 1.3 | 9.168 | 9.199 | 9.11 | [13 empty] | PCB 95+121 | 0.056 | <0.100 | 0.092 | 0.094 | 0.085
PCB 99 | [11 empty] | 26.6641 | 3.8772 | PCB 99 | 25.6 + 2.3 | 25.9 | 3.93 ± 0.17 | 4.32 | 3.50 | PCB 99 | 25.6 + 2.3 | 25.4 |  | 26 | 3.93 ± 0.17 | 3.77 |  | 3.96 | PCB 99 | 25.6 + 2.3 | 26.4 | 3.93 ± 0.17 | 3.95 | PCB 99 |  | 3.93 ± 0.17 |  | 5.03 | PCB 99 | 0.0418 |  | 3.93 ± 0.17 |  | 3.9 | 4.08 | 4.08 |  | 4.02 | 0.1 | 2.5 | [7 empty] | PCB 99 | 3.93 ± 0.17 | 4.021 | 3.821 | 3.846 | 25.6 + 2.3 | 26.324 | 26.117 | 26.796 | [13 empty] | PCB 99 | 0.116 | 3.93 ± 0.17 | 3.9 | 3.811 | 3.894
PCB 101 | [11 empty] | 36.5141 | 0.32813 | PCB 101 | 34.6 + 2.6 | 34.4 | 0.835 ± 0.7 | 0.340 | <0.100 | PCB 101 | 34.6 + 2.6 | 32.2 |  | 31.8 | 0.835 ± 0.7 | 0.295 |  | 0.274 | PCB 101 | 34.6 + 2.6 | 36.3 | 0.835 ± 0.7 | 0.222 | PCB 101 |  | 0.835 ± 0.7 |  | 0.261 | PCB 101 | 0.142 |  | 0.835 ± 0.7 |  | 0.416 | 0.432 | 0.411 |  | 0.42 | 0.011 | 2.5 | [7 empty] | PCB 101 | 0.835 ± 0.7 | 0.305 | 0.296 | 0.297 | 34.6 + 2.6 | 33.225 | 33.627 | 34.862 | [13 empty] | PCB 101 | 0.068 | 0.835 ± 0.7 | 0.311 | 0.308 | 0.276
PCB 105 |  | PCB 105 | 22.7 | 21.8 | 20.2 | 21.7 | 21.3 | 21.54 | 0.90719 | 19.9 + 0.93 |  | 20.991 | 1.6757 | PCB 105 | 19.9 + 0.93 | 18.4 | 1.84 ± 0.47 | 1.85 | 1.79 | PCB 105 | 19.9 + 0.93 | 19 |  | 19.2 | 1.84 ± 0.47 | 1.62 |  | 1.59 | PCB 105 | 19.9 + 0.93 | 21.2 | 1.84 ± 0.47 | 1.91 | PCB 105 |  | 1.84 ± 0.47 |  | 1.86 | PCB 105 | 0.152 |  | 1.84 ± 0.47 |  | 1.7 | 1.73 | 1.73 |  | 1.72 | 0.016 | 0.95 | [7 empty] | PCB 105 | 1.84 ± 0.47 | 1.851 | 1.838 | 1.868 | 19.9 + 0.93 | 20.208 | 20.549 | 20.931 | [13 empty] | PCB 105 | 0.105 | 1.84 ± 0.47 | 1.673 | 1.656 | 1.637
[20 empty] | PCB 106 |  | 0.0165 |  | 0.00915 | <0.100 | <LOD |  | 0.00624 | [5 empty] | PCB 106 |  | <0.100 |  | 0.0279 | PCB 106 | 0.0169 |  | <0.100 |  | 0.017 | 0.0188 | 0.0165 |  | 0.0175 | 0.0012 | 7 | [7 empty] | PCB 106 | <0.100 | 0.039 | 0.038 | 0.037 |  | 0.121 | 0.123 | 0.125 | [13 empty] | PCB 106 | 0.02 | <0.100 | 0.047 | 0.054 | 0.052
PCB 107 | [11 empty] | 9.0206 | 0.42153 | PCB 107 | 8.86 + 0.20 | 8.28 | <0.7; 0.472 ± 0.022 | 0.520 | 0.371 | PCB 107 | 8.86 + 0.20 | 7.23 |  | 6.96 | <0.7; 0.472 ± 0.022 | 0.373 |  | 0.369 | PCB 107 | 8.86 + 0.20 | 10.7 | <0.7; 0.472 ± 0.022 | 0.480 | PCB 107 |  | 0.472 ± 0.022 |  | 0.478 | PCB 107 | 0.0493 |  | 0.472 ± 0.022 |  | 0.425 | 0.417 | 0.401 |  | 0.414 | 0.012 | 3 | [7 empty] | PCB 107 | 0.472 ± 0.022 | 0.474 | 0.49 | 0.475 | 8.86 + 0.20 | 8.978 | 8.65 | 8.804 | [13 empty] | PCB 107 | 0.18 | 0.472 ± 0.022 | 0.452 | 0.458 | 0.455
PCB 110 | [11 empty] | 24.2975 | 0.06624 | PCB 110 | 22.8 + 2.0 | 26.0 | <0.100 | <0.100 | <0.100 | PCB 110 | 22.8 + 2.0 | 22.4 |  | 22.8 | <0.100 | 0.0747 |  | 0.0551 | PCB 110 | 22.8 + 2.0 | 26.8 | <0.100 | <0.100 | PCB 110 |  | <0.100 |  | <LOD | PCB 110 | 0.0278 |  | <0.100 |  | <LOD | 0.0351 | 0.0383 |  | 0.0367 | 0.0023 | 6.2 | [7 empty] | PCB 110 | <0.100 | 0.054 | 0.056 | 0.057 | 22.8 + 2.0 | 22.326 | 22.412 | 23.362 | [13 empty] | PCB 110 | 0.077 | <0.100 | 0.027 | 0.03 | 0.028
[20 empty] | PCB 112 |  | 0.0758 |  | 0.0858 |  | 0.0109 |  | 0.0141 | [5 empty] | PCB 112 |  |  |  | <LOD | PCB 112 | 0.00692 |  |  |  | 0.0183 | 0.0209 | 0.0196 |  | 0.0196 | 0.0013 | 6.6 | [7 empty] | PCB 112 |  | 0.028 | 0.029 | 0.03 |  | 0.121 | 0.12 | 0.12
PCB 114 | [11 empty] | 1.7227 | 0.25883 | PCB 114 |  | 0.656 | <0.5; 0.271 ± 0.017 | <0.100 | <0.100 | PCB 114 |  | 1.42 |  | 1.56 | <0.5; 0.271 ± 0.017 | 0.237 |  | 0.281 | PCB 114 |  | 1.29 | <0.5; 0.271 ± 0.017 | 0.368 | PCB 114 |  | 0.271 ± 0.017 |  | 0.254 | PCB 114 | 0.0338 |  | 0.271 ± 0.017 |  | 0.184 | 0.186 | 0.188 |  | 0.186 | 0.0019 | 1 | [7 empty] | PCB 114 | 0.271 ± 0.017 | 0.149 | 0.143 | 0.149 |  | 1.195 | 1.157 | 1.188 | [13 empty] | PCB 114 | 0.046 | 0.271 ± 0.017 | 0.264 | 0.255 | 0.267
PCB 118 |  | PCB 118 | 58.1 | 54.5 | 54.4 | 53.1 | 55.8 | 55.18 | 1.8913 | 52.1 + 1.0 |  | 52.2534 | 5.4309 | PCB 118+106 | 52.1 + 1.0; | 57.8 | 6.53 ± 1.1; | 6.50 | 5.84 | PCB 118 | 52.1 + 1.0 | 48.9 |  | 47.6 | 6.53 ± 1.1 | 5.51 |  | 5.33 | PCB 118 | 52.1 + 1.0 | 53.9 | 6.53 ± 1.1 | 5.84 | PCB 118 |  | 6.53 ± 1.1 |  | 7.28 | PCB 118 | 0.497 |  | 6.53 ± 1.1 |  | 5.99 | 5.88 | 5.81 |  | 5.89 | 0.091 | 1.6 | [7 empty] | PCB 118 | 6.53 ± 1.1 | 6.166 | 6.088 | 6.051 | 52.1 + 1.0 | 52.589 | 52.416 | 53.212 | [13 empty] | PCB 118 | 0.19 | 6.53 ± 1.1 | 5.735 | 5.76 | 5.637
PCB 119 | [11 empty] | 0.16832 | 0.05 | PCB 119 |  | 1.86 | <0.100 | <0.100 | <0.100 | PCB 119 |  | 0.852 |  | 0.952 | <0.100 | 0.0919 |  | 0.0971 | PCB 119 |  | 1.46 | <0.100 | <0.100 | PCB 119 |  | <0.100 |  | 0.085 | PCB 119 | 0.0101 |  | <0.100 |  | 0.0936 | 0.0818 | 0.0856 |  | 0.087 | 0.006 | 6.9 | [7 empty] | PCB 119 | <0.100 | 0.092 | 0.091 | 0.094 |  | 0.947 | 1.024 | 0.974 | [13 empty] | PCB 119 | 0.01 | <0.100 | <LOQ | 0.008 | 0.009
[20 empty] | PCB 127 |  | <LOD |  | <LOD | <0.100 | <LOD |  | <LOD | [5 empty] | PCB 127 |  | <0.100 |  | <LOD | PCB 127 | 0.0293 |  | <0.100 |  | <LOQ | <LOQ | <LOQ |  | <LOQ | [9 empty] | PCB 127 | <0.100 | 0.064 | 0.065 | 0.068 |  | 0.103 | 0.109 | 0.112
PCB 128 |  | PCB 128 | 25.6 | 22.8 | 23 | 23.6 | 28.9 | 24.78 | 2.5558 | 22.8 + 1.9 |  | 22.7802 | 0.60653 | [6 empty] | PCB 128 | 22.8 + 1.9 | 21.4 |  | 22.4 | 0.809 ± 0.077 | 0.709 |  | 0.724 | PCB 128 | 22.8 + 1.9 | 24.0 | 0.809 ± 0.077 | 0.814 | PCB 128 |  | 0.809 ± 0.077 |  | 0.895 | PCB 128 | 0.085 |  | 0.809 ± 0.077 |  | 0.742 | 0.764 | 0.741 |  | 0.749 | 0.013 | 1.7 | [7 empty] | PCB 128 | 0.809 ± 0.077 | 0.81 | 0.808 | 0.815 | 22.8 + 1.9 | 22.712 | 22.152 | 22.467 | [13 empty] | PCB 128 | 0.049 | 0.809 ± 0.077 | 0.77 | 0.804 | 0.778
PCB 130 | [11 empty] | 7.2117 | 0.37751 | PCB 130 |  | 5.95 | 0.298 ± 0.018 | 0.310 | <0.100 | PCB 130 |  | 6.81 |  | 6.68 | 0.298 ± 0.018 | 0.305 |  | 0.313 | PCB 130 |  | 8.08 | 0.298 ± 0.018 | 0.350 | PCB 130 |  | 0.298 ± 0.018 |  | 0.362 | PCB 130 | 0.0386 |  | 0.298 ± 0.018 |  | 0.291 | 0.311 | 0.29 |  | 0.297 | 0.012 | 3.9 | [7 empty] | PCB 130 | 0.298 ± 0.018 | 0.297 | 0.298 | 0.298 |  | 6.981 | 6.901 | 7.053 | [13 empty] | PCB 130 | 0.038 | 0.298 ± 0.018 | 0.3 | 0.323 | 0.299
[14 empty] | PCB 132 | 5.83 + 0.76 | 6.05 |  | 0.121 | <0.100
PCB 137 | [11 empty] | 5.392 | 0.29694 | PCB 137 |  | 4.42 | <4; 1.12 ± 2.2 | 0.313 | 0.105 | PCB 137 |  | 4.43 |  | 4.41 | <4 | 0.187 |  | 0.197 | PCB 137 |  | 6.12 | <4; 1.12 ± 2.2 | 0.981 | PCB 137 |  | 1.12 ± 2.2 |  | 0.28 | PCB 137 | 0.035 |  | 1.12 ± 2.2 |  | 0.207 | 0.208 | 0.206 |  | 0.207 | 0.00095 | 0.46 | [7 empty] | PCB 137 | 1.12 ± 2.2 | 0.205 | 0.209 | 0.205 |  | 4.663 | 4.699 | 4.738 | [13 empty] | PCB 137 | 0.008 | 1.12 ± 2.2 | 0.243 | 0.234 | 0.229
PCB 138 | [11 empty] | 126.8835 | 4.5964 | PCB 138 | 115 + 13 | 115 | 5.76 ± 1.5 | 5.78 | 4.85 | PCB 138 | 115 + 13 | 104 |  | 103 | 5.76 ± 1.5 | 4.51 |  | 4.43 | PCB 138 | 115 + 13 | 118 | 5.76 ± 1.5 | 4.93 | PCB 138 |  | 5.76 ± 1.5 |  | 5.22 | PCB 138 | 0.753 |  | 5.76 ± 1.5 |  | 4.98 | 4.96 | 4.69 |  | 4.88 | 0.16 | 3.3 | [7 empty] | PCB 138 | 5.76 ± 1.5 | 5.072 | 5.085 | 4.936 | 115 + 13 | 107.124 | 108.34 | 105.46 | [13 empty] | PCB 138 | 0.286 | 5.76 ± 1.5 | 4.704 | 4.724 | 4.792
PCB 146 | [11 empty] | 31.6197 | 2.3774 | PCB 146 | 30.1 + 3.5 | 29.3 | 2.47 ± 0.022 | 2.33 | 2.03 | PCB 146 | 30.1 + 3.5 | 26.9 |  | 26.8 | 2.47 ± 0.022 | 2.06 |  | 2.01 | PCB 146 | 30.1 + 3.5 | 39.8 | 2.47 ± 0.022 | 2.81 | PCB 146 |  | 2.47 ± 0.022 |  | 2.74 | PCB 146 | 0.212 |  | 2.47 ± 0.022 |  | 2.21 | 2.22 | 2.2 |  | 2.21 | 0.0094 | 0.42 | [7 empty] | PCB 146 | 2.47 ± 0.022 | 2.329 | 2.278 | 2.314 | 30.1 + 3.5 | 28.142 | 28.151 | 29.357 | [13 empty] | PCB 146 | 0.203 | 2.47 ± 0.022 | 2.453 | 2.478 | 2.485
PCB 149 | [11 empty] | 27.3431 | 0.27082 | PCB 149 | 26.3 + 1.3 | 30.5 | 0.403 ± 0.1 | 0.439 | 0.322 | PCB 149 | 26.3 + 1.3 | 25.4 |  | 25.1 | 0.403 ± 0.1 | 0.221 |  | 0.264 | PCB 149 | 26.3 + 1.3 | 27.2 | 0.403 ± 0.1 | 0.298 | PCB 149 |  | 0.403 ± 0.1 |  | 0.45 | PCB 149 | 0.2 |  | 0.403 ± 0.1 |  | 0.407 | 0.429 | 0.391 |  | 0.409 | 0.019 | 4.7 | [7 empty] | PCB 149 | 0.403 ± 0.1 | 0.446 | 0.439 | 0.425 | 26.3 + 1.3 | 27.767 | 26.402 | 27.204 | [13 empty] | PCB 149 | 0.085 | 0.403 ± 0.1 | 0.314 | 0.304 | 0.304
PCB 151 | [11 empty] | 9.7515 | 0.054433 | PCB 151 |  | 9.40 | <0.2; 0.115 ± 0.0090 | <0.100 | <0.100 | PCB 151 |  | 8.34 |  | 8.4 | <0.2 | <LOQ |  | <LOQ | PCB 151 |  | 14.2 | <0.2; 0.115 ± 0.0090 | <0.100 | PCB 151 |  | <0.100 |  | 0.0812 | PCB 151 | 0.0155 |  | <0.100 |  | <LOQ | 0.0149 | 0.0154 |  | 0.0152 | 0.00036 | 2.3 | [7 empty] | PCB 151 | <0.100 | <LOD | <LOD | 0.026 |  | 7.745 | 7.857 | 8.273 | [13 empty] | PCB 151 | 0.039 | <0.100 | 0.015 | 0.017 | 0.015
PCB 153+132 |  | PCB 153 | 184 | 175 | 178 | 179 | 171 | 177.4 | 4.827 | 169.8 + 8.9 |  | 175.2542 | 10.2417 | PCB 153 | 170 + 9 | 164 | 11.0 ± 0.26 | 14.6 | 16.0 | PCB 153+132 | 170 + 9 | 167 |  | 167 | 11.0 ± 0.26 | 10 |  | 9.98 | PCB 153+132 | 170 + 9; 5.83 + 0.76 | 175 | 11.0 ± 0.26 | 10.7 | PCB 153+132 |  | 11.0 ± 0.26 |  | 12.8 | PCB 153+132 | 1.26 |  | 11.0 ± 0.26 |  | 11.2 | 11.2 | 11.1 |  | 11.1 | 0.045 | 0.41 | [7 empty] | PCB 153+132 | 11.0 ± 0.26 | 10.577 | 10.4 | 10.424 | 170 + 9; 5.83 + 0.76 | 171.939 | 172.759 | 176.005 | [13 empty] | PCB 153+132 | 0.516 | 11.0 ± 0.26 | 10.82 | 10.9 | 10.88
PCB 154 | [11 empty] | 1.5944 | 0.13086 | PCB 154 |  | 1.59 | <0.2; 0.138 ± 0.011 | <0.100 | <0.100 | PCB 154 |  | 1.29 |  | 1.26 | <0.2 | 0.05 |  | 0.0573 | PCB 154 |  | 2.04 | <0.2; 0.138 ± 0.011 | 0.456 | PCB 154 |  | <0.100 |  | 0.0994 | PCB 154 | 0.0601 |  | <0.100 |  | 0.0947 | 0.0952 | 0.0937 |  | 0.0945 | 0.00078 | 0.82 | [7 empty] | PCB 154 | <0.100 | 0.081 | 0.08 | 0.081 |  | 1.169 | 1.153 | 1.222 | [13 empty] | PCB 154 | 0.022 | <0.100 | 0.038 | 0.035 | 0.032
PCB 156 | [11 empty] | 9.4186 | 0.51762 | PCB 156 | 9.52 + 0.51 | 9.55 | 0.557 ± 0.12 | 0.512 | 0.306 | PCB 156 | 9.52 + 0.51 | 7.89 |  | 7.93 | 0.557 ± 0.12 | 0.473 |  | 0.438 | PCB 156 | 9.52 + 0.51 | 7.42 | 0.557 ± 0.12 | 0.574 | PCB 156 |  | 0.557 ± 0.12 |  | 0.694 | PCB 156 | 0.0848 |  | 0.557 ± 0.12 |  | 0.451 | 0.467 | 0.465 |  | 0.461 | 0.0086 | 1.9 | [7 empty] | PCB 156 | 0.557 ± 0.12 | 0.606 | 0.588 | 0.593 | 9.52 + 0.51 | 9.826 | 9.713 | 9.723 | [13 empty] | PCB 156 | 0.06 | 0.557 ± 0.12 | 0.447 | 0.478 | 0.439
PCB 157 | [11 empty] | 3.2431 | 0.13515 | PCB 157 |  | 1.89 | 0.210 ± 0.014 | 0.253 | 0.118 | PCB 157 |  | 2.06 |  | 1.95 | 0.210 ± 0.014 | 0.13 |  | 0.117 | PCB 157 |  | 3.04 | 0.210 ± 0.014 | 0.188 | PCB 157 |  | 0.210 ± 0.014 |  | 0.195 | PCB 157 | 0.0354 |  | 0.210 ± 0.014 |  | 0.128 | 0.129 | 0.13 |  | 0.129 | 0.0011 | 0.82 | [7 empty] | PCB 157 | 0.210 ± 0.014 | 0.139 | 0.149 | 0.141 |  | 1.857 | 1.845 | 1.879 | [13 empty] | PCB 157 | 0.043 | 0.210 ± 0.014 | 0.156 | 0.166 | 0.152
PCB 158 | [11 empty] | 5.9054 | 0.15383 | PCB 158 | 7.66 + 0.88 | 7.01 | 0.317 ± 0.063 | 0.248 | 0.333 | PCB 158 | 7.66 + 0.88 | 7.8 |  | 7.99 | 0.317 ± 0.063 | 0.297 |  | 0.299 | PCB 158 | 7.66 + 0.88 | 6.62 | 0.317 ± 0.063 | 0.279 | PCB 158 |  | 0.317 ± 0.063 |  | 0.344 | PCB 158 | 0.0562 |  | 0.317 ± 0.063 |  | 0.216 | 0.237 | 0.211 |  | 0.221 | 0.014 | 6.3 | [7 empty] | PCB 158 | 0.317 ± 0.063 | 0.297 | 0.308 | 0.314 | 7.66 + 0.88 | 7.428 | 7.681 | 7.736 | [13 empty] | PCB 158 | 0.013 | 0.317 ± 0.063 | 0.284 | 0.298 | 0.285
PCB 159 | [11 empty] | 0.12662 | 0.05 | PCB 159 |  | 0.137 | <0.100 | 0.159 | <0.100 | PCB 159 |  | 0.146 |  | 0.134 | <0.100 | <LOQ |  | <LOQ | PCB 159 |  | 0.109 | <0.100 | <0.100 | PCB 159 |  | <0.100 |  | <LOQ | PCB 159 | 0.00257 |  | <0.100 |  | <LOQ | 0.00343 | 0.00295 |  | 0.00319 | 0.00034 | 11 | [7 empty] | PCB 159 | <0.100 | 0.014 | 0.016 | 0.015 |  | 0.205 | 0.195 | 0.201 | [13 empty] | PCB 159 | 0.006 | <0.100 | 0.019 | 0.018 | 0.017
PCB 163 | [11 empty] | 33.7098 | 1.8962 | PCB 163 | 31.8 + 0.8 | 30.2 | 2.00 ± 0.16 | 2.28 | 2.12 | PCB 163 | 31.8 + 0.8 | 31.9 |  | 32.6 | 2.00 ± 0.16 | 1.98 |  | 1.92 | PCB 163 | 31.8 + 0.8 | 36.2 | 2.00 ± 0.16 | 2.04 | PCB 163 |  | 2.00 ± 0.16 |  | 2.41 | PCB 163 | 0.205 |  | 2.00 ± 0.16 |  | 1.95 | 1.93 | 1.94 |  | 1.94 | 0.0098 | 0.5 | [7 empty] | PCB 163 | 2.00 ± 0.16 | 2.004 | 1.984 | 1.991 | 31.8 + 0.8 | 32.343 | 31.478 | 32.367 | [13 empty] | PCB 163 | 0.078 | 2.00 ± 0.16 | 2.02 | 2.024 | 2.128
PCB 165 | [11 empty] | 0.35251 | 0.099462 | PCB 165 |  | 0.416 | 0.117 ± 0.0090 | 0.122 | <0.100 | PCB 165 |  | 0.229 |  | 0.232 | 0.117 ± 0.0090 | 0.0121 |  | 0.00662 | PCB 165 |  | 0.210 | 0.117 ± 0.0090 | <0.100 | PCB 165 |  | 0.117 ± 0.0090 |  | 0.0476 | PCB 165 | 0.036 |  | 0.117 ± 0.0090 |  | 0.0428 | 0.0406 | 0.0385 |  | 0.0406 | 0.0022 | 5.3 | [7 empty] | PCB 165 | 0.117 ± 0.0090 | 0.037 | 0.04 | 0.037 |  | 0.205 | 0.213 | 0.208 | [13 empty] | PCB 165 | 0.046 | 0.117 ± 0.0090 | 0.049 | 0.057 | 0.053
PCB 166 | [11 empty] | 0.90643 | 0.32238 | PCB 166 |  | 0.945 | 0.253 ± 0.023 | 0.187 | <0.100 | PCB 166 |  | 0.513 |  | 0.513 | 0.253 ± 0.023 | 0.0407 |  | 0.0589 | PCB 166 |  | 1.25 | 0.253 ± 0.023 | 0.882 | PCB 166 |  | 0.253 ± 0.023 |  | 0.0602 | PCB 166 | 0.0281 |  | 0.253 ± 0.023 |  | 0.123 | 0.113 | 0.116 |  | 0.117 | 0.005 | 4.3 | [7 empty] | PCB 166 | 0.253 ± 0.023 | 0.077 | 0.079 | 0.083 |  | 0.561 | 0.573 | 0.6 | [13 empty] | PCB 166 | 0.047 | 0.253 ± 0.023 | 0.056 | 0.052 | 0.049
PCB 167 | [11 empty] | 5.6896 | 0.5071 | PCB 167+128 | ; 22.8 + 1.9 | 34.1 | 0.468 ± 0.037; 0.809 ± 0.077 | 1.11 | 0.251 | PCB 167 |  | 5.08 |  | 4.86 | 0.468 ± 0.037 | 0.328 |  | 0.31 | PCB 167 |  | 6.28 | 0.468 ± 0.037 | 0.396 | PCB 167 |  | 0.468 ± 0.037 |  | 0.361 | PCB 167 | 0.0542 |  | 0.468 ± 0.037 |  | 0.371 | 0.36 | 0.347 |  | 0.359 | 0.012 | 3.4 | [7 empty] | PCB 167 | 0.468 ± 0.037 | 0.391 | 0.406 | 0.409 |  | 4.961 | 5.076 | 5.062 | [13 empty] | PCB 167 | 0.017 | 0.468 ± 0.037 | 0.365 | 0.353 | 0.343
PCB 170 |  | PCB 170 | 27 | 26.2 | 24 | 26.7 | 28.8 | 26.54 | 1.7257 | 25.2 + 2.2 |  | 27.9768 | 0.79732 | PCB 170 | 25.2 + 2.2 | 28.3 | 1.47 ± 0.72 | 1.27 | 0.878 | PCB 170 | 25.2 + 2.2 | 27 |  | 28.9 | 1.47 ± 0.72 | 0.863 |  | 0.831 | PCB 170 | 25.2 + 2.2 | 28.1 | 1.47 ± 0.72 | 1.02 | PCB 170 |  | 1.47 ± 0.72 |  | 1.25 | PCB 170 | 0.245 |  | 1.47 ± 0.72 |  | 1.02 | 1.04 | 0.962 |  | 1.01 | 0.039 | 3.9 | [7 empty] | PCB 170 | 1.47 ± 0.72 | 0.993 | 1.091 | 1.046 | 25.2 + 2.2 | 26.054 | 26.744 | 26.801 | [13 empty] | PCB 170 | 0.061 | 1.47 ± 0.72 | 1.02 | 0.954 | 0.961
PCB 172 | [11 empty] | 7.963 | 0.43602 | PCB 172 |  | 6.44 | 0.371 ± 0.028 | 0.402 | 0.291 | PCB 172 |  | 7.18 |  | 6.8 | 0.371 ± 0.028 | 0.323 |  | 0.309 | PCB 172 |  | 9.39 | 0.371 ± 0.028 | 0.383 | PCB 172 |  | 0.371 ± 0.028 |  | 0.397 | PCB 172 | 0.0505 |  | 0.371 ± 0.028 |  | 0.315 | 0.311 | 0.312 |  | 0.313 | 0.0021 | 0.67 | [7 empty] | PCB 172 | 0.371 ± 0.028 | 0.367 | 0.371 | 0.356 |  | 6.282 | 6.375 | 6.195 | [13 empty] | PCB 172 | 0.067 | 0.371 ± 0.028 | 0.377 | 0.357 | 0.376
PCB 174 | [11 empty] | 9.6006 | 0.091671 | PCB 174 | 9.3 + 1.3 | 9.89 | <0.100 | <0.100 | <0.100 | PCB 174 | 9.3 + 1.3 | 8.44 |  | 7.96 | <0.100 | 0.0284 |  | 0.0286 | PCB 174 | 9.3 + 1.3 | 11.7 | <0.100 | <0.100 | PCB 174 |  | <0.100 |  | 0.0831 | PCB 174 | 0.0693 |  | <0.100 |  | 0.0923 | 0.0914 | 0.093 |  | 0.0922 | 0.00078 | 0.85 | [7 empty] | PCB 174 | <0.100 | 0.084 | 0.09 | 0.086 | 9.3 + 1.3 | 10.244 | 10.269 | 10.477 | [13 empty] | PCB 174 | 0.05 | <0.100 | 0.056 | 0.053 | 0.052
PCB 175 | [11 empty] | 1.8903 | 0.24904 | PCB 175 |  | 2.07 | 0.266 ± 0.019 | 0.235 | <0.100 | PCB 175 |  | 1.55 |  | 1.43 | 0.266 ± 0.019 | <LOQ |  | 0.0818 | PCB 175 |  | 1.89 | 0.266 ± 0.019 | 0.165 | PCB 175 |  | 0.266 ± 0.019 |  | 0.0669 | PCB 175 | 0.035 |  | 0.266 ± 0.019 |  | 0.0903 | 0.0884 | 0.083 |  | 0.0872 | 0.0037 | 4.3 | [7 empty] | PCB 175 | 0.266 ± 0.019 | 0.075 | 0.079 | 0.074 |  | 1.36 | 1.331 | 1.353 | [13 empty] | PCB 175 | 0.05 | 0.266 ± 0.019 | 0.122 | 0.121 | 0.132
PCB 176 | [11 empty] | 0.7897 | 0.1771 | PCB 176 |  | 1.10 | <0.2; 0.171 ± 0.024 | 0.128 | <0.100 | PCB 176 |  | 0.569 |  | 0.527 | <0.2 | <LOD |  | 0.00286 | PCB 176 |  | 0.516 | <0.2; 0.171 ± 0.024 | 0.135 | PCB 176 |  | 0.171 ± 0.024 |  | 0.0329 | PCB 176 | 0.0339 |  | 0.171 ± 0.024 |  | 0.0325 | 0.038 | 0.033 |  | 0.0345 | 0.003 | 8.8 | [7 empty] | PCB 176 | 0.171 ± 0.024 | 0.019 | 0.019 | 0.019 |  | 0.355 | 0.343 | 0.35 | [13 empty] | PCB 176 | 0.057 | 0.171 ± 0.024 | <LOQ | <LOQ | <LOQ
PCB 177 | [11 empty] | 15.0137 | 0.31962 | PCB 177 |  | 15.9 | 0.395 ± 0.092 | 0.402 | 0.204 | PCB 177 |  | 13.7 |  | 13.8 | 0.395 ± 0.092 | 0.251 |  | 0.275 | PCB 177 |  | 18.0 | 0.395 ± 0.092 | 0.379 | PCB 177 |  | 0.395 ± 0.092 |  | 0.304 | PCB 177 | 0.105 |  | 0.395 ± 0.092 |  | 0.352 | 0.347 | 0.322 |  | 0.34 | 0.016 | 4.8 | [7 empty] | PCB 177 | 0.395 ± 0.092 | 0.31 | 0.328 | 0.314 |  | 14.951 | 15.08 | 14.943 | [13 empty] | PCB 177 | 0.034 | 0.395 ± 0.092 | 0.304 | 0.303 | 0.312
PCB 178 | [11 empty] | 14.0889 | 0.45771 | PCB 178 |  | 13.3 | 0.429 ± 0.02 | 0.481 | 0.346 | PCB 178 |  | 9.85 |  | 9.98 | 0.429 ± 0.020 | 0.372 |  | 0.414 | PCB 178 |  | 14.4 | 0.429 ± 0.02 | 0.450 | PCB 178 |  | 0.418 ± 0.019 |  | 0.301 | PCB 178 | 0.0878 |  | 0.418 ± 0.019 |  | 0.431 | 0.425 | 0.416 |  | 0.424 | 0.0077 | 1.8 | [7 empty] | PCB 178 | 0.418 ± 0.019 | 0.411 | 0.424 | 0.429 |  | 9.652 | 9.542 | 9.659 | [13 empty] | PCB 178 | 0.061 | 0.418 ± 0.019 | 0.416 | 0.405 | 0.405
PCB 180+193 |  | PCB 180 | 67.1 | 68.3 | 76.3 | 76.7 | 73.4 | 72.36 | 4.4607 | 74.4 + 4.0 |  | 80.7554 | 2.115 | PCB 180+193 | 74.4 + 4.0; 5.78 + 0.72 | 77.1 | 2.28 ± 0.18 | 2.36 | 2.00 | PCB 180+193 | 74.4 + 4.0; 5.78 + 0.72 | 73.6 |  | 74.1 | 2.28 ± 0.18 | 1.9 |  | 1.97 | PCB 180+193 | 74.4 + 4.0; 5.78 + 0.72 | 89.1 | 2.28 ± 0.18 | 2.24 | PCB 180+193 |  | 2.28 ± 0.18 |  | 2.48 | PCB 180+193 | 0.632 |  | 2.28 ± 0.18 |  | 2.3 | 2.27 | 2.27 |  | 2.28 | 0.017 | 0.75 | [7 empty] | PCB 180+193 | 2.28 ± 0.18 | 2.404 | 2.411 | 2.325 | 74.4 + 4.0; 5.78 + 0.72 | 75.193 | 75.425 | 74.322 | [13 empty] | PCB 180+193 | 0.195 | 2.28 ± 0.18 | 2.231 | 2.213 | 2.186
PCB 183 | [11 empty] | 21.2745 | 0.78535 | PCB 183 | 21.9 + 2.5 | 21.1 | 0.838 ± 0.094 | 0.817 | 0.556 | PCB 183 | 21.9 + 2.5 | 20 |  | 19.7 | 0.838 ± 0.094 | 0.633 |  | 0.641 | PCB 183 | 21.9 + 2.5 | 22.2 | 0.838 ± 0.094 | 0.907 | PCB 183 |  | 0.838 ± 0.094 |  | 0.712 | PCB 183 | 0.156 |  | 0.838 ± 0.094 |  | 0.731 | 0.733 | 0.7 |  | 0.721 | 0.019 | 2.6 | [7 empty] | PCB 183 | 0.838 ± 0.094 | 0.831 | 0.809 | 0.797 | 21.9 + 2.5 | 19.596 | 19.528 | 19.627 | [13 empty] | PCB 183 | 0.045 | 0.838 ± 0.094 | 0.764 | 0.765 | 0.777
PCB 185 | [11 empty] | 2.1246 | 0.092716 | PCB 185 |  | 2.71 | <0.100 | <0.100 | <0.100 | PCB 185 |  | 1.9 |  | 1.97 | <0.100 | 0.0133 |  | <LOD | PCB 185 |  | 2.31 | <0.100 | <0.100 | PCB 185 |  | <0.100 |  | 0.0251 | PCB 185 | 0.0149 |  | <0.100 |  | 0.0223 | 0.0234 | 0.023 |  | 0.0229 | 0.00053 | 2.3 | [7 empty] | PCB 185 | <0.100 | <LOD | 0.027 | 0.025 |  | 1.475 | 1.505 | 1.487 | [13 empty] | PCB 185 | 0.019 | <0.100 | 0.021 | 0.019 | 0.021
PCB 187 |  | PCB 187 | 54.1 | 53.8 | 50 | 54.4 | 54.7 | 53.4 | 1.93 | 55.2 + 2.1 |  | 56.4728 | 2.7814 | PCB 187 | 55.2 + 2.1 | 55.6 | 2.94 ± 0.09 | 3.05 | 2.67 | PCB 187 | 55.2 + 2.1 | 54.7 |  | 55.1 | 2.94 ± 0.09 | 2.74 |  | 2.93 | PCB 187 | 55.2 + 2.1 | 54.8 | 2.94 ± 0.09 | 2.92 | PCB 187 |  | 2.94 ± 0.09 |  | 2.63 | PCB 187 | 0.151 |  | 2.94 ± 0.09 |  | 2.92 | 2.89 | 2.85 |  | 2.89 | 0.037 | 1.3 | [7 empty] | PCB 187 | 2.94 ± 0.09 | 2.964 | 3.029 | 2.964 | 55.2 + 2.1 | 55.565 | 55.164 | 55.083 | [13 empty] | PCB 187 | 0.15 | 2.94 ± 0.09 | 2.832 | 2.968 | 2.996
PCB 188 | [11 empty] | 0.52438 | 0.17258 | PCB 188 |  | 0.571 | <0.3; 0.194 ± 0.019 | 0.168 | <0.100 | PCB 188 |  | 0.301 |  | 0.305 | <0.3; 0.194 ± 0.019 | 0.015 |  | 0.0122 | PCB 188 |  | 0.320 | <0.3; 0.194 ± 0.019 | <0.100 | PCB 188 |  | 0.194 ± 0.019 |  | 0.037 | PCB 188 | 0.0197 |  | 0.194 ± 0.019 |  | 0.0366 | 0.0356 | 0.0328 |  | 0.035 | 0.002 | 5.7 | [7 empty] | PCB 188 | 0.194 ± 0.019 | 0.022 | 0.022 | 0.02 |  | 0.195 | 0.193 | 0.196 | [13 empty] | PCB 188 | 0.042 | 0.194 ± 0.019 | 0.023 | 0.027 | 0.025
PCB 189 | [11 empty] | 1.589 | 0.078753 | PCB 189 |  | 1.11 | <0.100 | <0.100 | <0.100 | PCB 189 |  | 1.29 |  | 1.33 | <0.100 | 0.0435 |  | 0.0385 | PCB 189 |  | 1.72 | <0.100 | <0.100 | PCB 189 |  | <0.100 |  | 0.0214 | PCB 189 | 0.0121 |  | <0.100 |  | <LOQ | <LOQ | 0.0446 |  | <LOQ | [9 empty] | PCB 189 | <0.100 | 0.055 | 0.059 | 0.06 |  | 2.277 | 2.421 | 2.354 | [13 empty] | PCB 189 | 0.004 | <0.100 | 0.062 | 0.061 | 0.064
PCB 191 | [11 empty] | 0.91249 | 0.05 | PCB 191 |  | 1.09 | <0.100 | <0.100 | <0.100 | PCB 191 |  | 0.863 |  | 0.781 | <0.100 | 0.0146 |  | 0.0087 | PCB 191 |  | 0.990 | <0.100 | <0.100 | PCB 191 |  | <0.100 |  | 0.0298 | PCB 191 | 0.027 |  | <0.100 |  | <LOQ | 0.0257 | <LOQ |  | <LOQ | [9 empty] | PCB 191 | <0.100 | 0.015 | 0.014 | 0.016 |  | 0.816 | 0.803 | 0.8 | [13 empty] | PCB 191 | 0.003 | <0.100 | 0.015 | 0.015 | 0.018
PCB 194 | [11 empty] | 14.0511 | 0.3337 | PCB 194 | 13.0 + 1.3 | 13.0 | <0.6; 0.316 ± 0.029 | 0.369 | 0.295 | PCB 194 | 13.0 + 1.3 | 12.4 |  | 12.3 | <0.6; 0.316 ± 0.029 | 0.297 |  | 0.314 | PCB 194 | 13.0 + 1.3 | 14.8 | <0.6; 0.316 ± 0.029 | 0.380 | PCB 194 |  | 0.316 ± 0.029 |  | 0.332 | PCB 194 | 0.0775 |  | 0.316 ± 0.029 |  | 0.303 | 0.313 | 0.325 |  | 0.314 | 0.011 | 3.6 | [7 empty] | PCB 194 | 0.316 ± 0.029 | 0.322 | 0.338 | 0.333 | 13.0 + 1.3 | 12.617 | 12.64 | 12.325 | [13 empty] | PCB 194 | 0.073 | 0.316 ± 0.029 | 0.288 | 0.298 | 0.292
PCB 195 |  | PCB 195 | 7.35 | 5.71 | 5.8 | 6.24 | 5.73 | 6.166 | 0.6963 | 5.30 + 0.45 |  | 4.7797 | 0.12273 | PCB 195 | 5.30 + 0.45 | 5.15 | <0.3; 0.148 ± 0.019 | 0.786 | 0.563 | PCB 195 | 5.30 + 0.45 | 4.48 |  | 5.14 | <0.3; 0.148 ± 0.019 | 0.105 |  | 0.129 | PCB 195 | 5.30 + 0.45 | 4.96 | <0.3; 0.148 ± 0.019 | 0.874 | PCB 195 |  | 0.148 ± 0.019 |  | 0.139 | PCB 195 | 0.0506 |  | 0.148 ± 0.019 |  | 0.128 | 0.132 | 0.131 |  | 0.131 | 0.0023 | 1.8 | [7 empty] | PCB 195 | 0.148 ± 0.019 | 0.142 | 0.143 | 0.147 | 5.30 + 0.45 | 5.674 | 5.468 | 5.698 | [13 empty] | PCB 195 | 0.01 | 0.148 ± 0.019 | 0.148 | 0.136 | 0.131
PCB 196+203 | [11 empty] | 19.7712 | 0.70218 | PCB 196+203 |  | 52.9 | <1; 0.724 ± 0.061 | 3.14 | 18.3 | PCB 196 |  | 19.4 |  | 19 | <1; 0.724 ± 0.061 | 0.408 |  | 0.394 | PCB 196+203 |  | 18.9 | <1; 0.724 ± 0.061 | 1.02 | PCB 196 |  | 0.724 ± 0.061 |  | 0.758 | PCB 196 | 0.336 |  | 0.724 ± 0.061 |  | 0.546 | 0.587 | 0.541 |  | 0.558 | 0.025 | 4.5 | [7 empty] | PCB 196 | 0.724 ± 0.061 | 0.78 | 0.765 | 0.736 |  | 23.254 | 23.607 | 23.905 | [13 empty] | PCB 196 | 0.089 | 0.724 ± 0.061 | 0.499 | 0.465 | 0.467
PCB 197 | [11 empty] | 1.2862 | 0.11835 | PCB 197 |  | 2.05 | 0.148 ± 0.011 | 0.131 | 0.135 | PCB 197 |  | 1.37 |  | 1.38 | 0.148 ± 0.011 | 0.0473 |  | 0.0495 | PCB 197 |  | 1.17 | 0.148 ± 0.011 | 0.111 | PCB 197 |  | 0.148 ± 0.011 |  | 0.0654 | PCB 197 | 0.0265 |  | 0.148 ± 0.011 |  | 0.0473 | 0.0488 | 0.0515 |  | 0.0492 | 0.0021 | 4.4 | [7 empty] | PCB 197 | 0.148 ± 0.011 | 0.043 | 0.044 | 0.043 |  | 0.932 | 0.908 | 0.934 | [13 empty] | PCB 197 | 0.032 | 0.148 ± 0.011 | 0.071 | 0.068 | 0.072
PCB 199 | [11 empty] | 16.3166 | 0.65213 | PCB 199 |  | 24.0 | <1.5; 0.631 ± 0.022 | 0.521 | 0.738 | PCB 199 |  | 60.2 |  | 60.9 | <1.5; 0.631 ± 0.022 | 2.11 |  | 2.15 | PCB 199 |  | 16.8 | <1.5; 0.631 ± 0.022 | 0.667 | PCB 199 |  | 0.631 ± 0.022 |  | 2.99 | PCB 199 | 0.543 |  | 0.631 ± 0.022 |  | 2 | 2.06 | 2.14 |  | 2.07 | 0.071 | 3.4 | [7 empty] | PCB 199 | 0.631 ± 0.022 | 2.116 | 2.208 | 2.097 |  | 49.947 | 51.104 | 51.954 | [13 empty] | PCB 199 | 0.083 | 0.631 ± 0.022 | 0.632 | 0.648 | 0.624
PCB 200 | [11 empty] | 0.2688 | 0.05 | PCB 200 |  | 0.333 | <0.3 | <0.100 | <0.100 | PCB 200 |  | 0.00815 |  | 0.00786 | <0.3 | <LOD |  | <LOD | PCB 200 |  | <0.100 | <0.3 | <0.100 | PCB 200 |  | <0.100 |  | <LOD | PCB 200 | 0.0222 |  | <0.100 |  | <LOQ | <LOQ | <LOQ |  | <LOQ | [9 empty] | PCB 200 | <0.100 | <LOD | <LOD | <LOD |  | 0.011 | 0.011 | 0.012 | [13 empty] | PCB 200 | 0.095 | <0.100 | <LOD | <LOD | <LOD
PCB 201 | [11 empty] | 2.8431 | 0.13229 | PCB 201 | 2.83 + 0.13 | 3.01 | <0.3; 0.198 ± 0.017 | 0.169 | 0.189 | PCB 201 | 2.83 + 0.13 | 13.9 |  | 12.8 | <0.3; 0.198 ± 0.017 | 0.526 |  | 0.489 | PCB 201 | 2.83 + 0.13 | 2.55 | <0.3; 0.198 ± 0.017 | 0.205 | PCB 201 |  | 0.198 ± 0.017 |  | 0.948 | PCB 201 | 0.207 |  | 0.198 ± 0.017 |  | 0.841 | 0.815 | 0.859 |  | 0.839 | 0.022 | 2.7 | [7 empty] | PCB 201 | 0.198 ± 0.017 | 0.845 | 0.818 | 0.782 | 2.83 + 0.13 | 14.357 | 14.275 | 14.093 | [13 empty] | PCB 201 | 0.022 | 0.198 ± 0.017 | 0.139 | 0.134 | 0.137
PCB 202 | [11 empty] | 5.758 | 0.068541 | [6 empty] | PCB 202 |  | 4.13 |  | 4.14 | <0.100 | 0.0408 |  | 0.0271 | PCB 202 |  | 4.53 | <0.100 | <0.100 | PCB 202 |  | <0.100 |  | 0.0564 | PCB 202 | 0.147 |  | <0.100 |  | <LOQ | <LOQ | <LOQ |  | <LOQ | [9 empty] | PCB 202 | <0.100 | <LOD | <LOD | <LOD |  | 2.999 | 3.053 | 3.103 | [13 empty] | PCB 202 | 0.045 | <0.100 | 0.042 | 0.038 | 0.041
PCB 206 |  | PCB 206 | 5.58 | 4.96 | 5.54 | 5.5 | 5.15 | 5.346 | 0.27546 | 5.40 + 0.43 |  | 5.5992 | 0.085847 | PCB 206 | 5.40 + 0.43 | 5.22 | <0.100 | 0.673 | <0.100 | PCB 206 | 5.40 + 0.43 | 4.56 |  | 4.57 | <0.100 | 0.0985 |  | 0.0633 | PCB 206 | 5.40 + 0.43 | 6.55 | <0.100 | 0.132 | PCB 206 |  | <0.100 |  | 0.169 | PCB 206 | 0.238 |  | <0.100 |  | <LOQ | 0.221 | <LOQ |  | <LOQ | [9 empty] | PCB 206 | <0.100 | 0.099 | 0.096 | 0.098 | 5.40 + 0.43 | 4.611 | 4.616 | 4.89 | [13 empty] | PCB 206 | 0.027 | <0.100 | 0.083 | 0.078 | 0.085
PCB 207 | [11 empty] | 1.0858 | 0.05 | PCB 207 |  | 0.833 | <0.100 | <0.100 | <0.100 | PCB 207 |  | 0.991 |  | 0.84 | <0.100 | 0.0436 |  | 0.0325 | PCB 207 |  | 1.48 | <0.100 | <0.100 | PCB 207 |  | <0.100 |  | 0.0725 | PCB 207 | 0.0424 |  | <0.100 |  | <LOQ | 0.087 | 0.0818 |  | 0.0844 | 0.0037 | 4.4 | [7 empty] | PCB 207 | <0.100 | 0.074 | 0.074 | 0.079 |  | 0.914 | 0.935 | 0.963 | [13 empty] | PCB 207 | 0.027 | <0.100 | 0.036 | 0.037 | 0.04
PCB 208 | [11 empty] | 1.511 | 0.0077568 | PCB 208 |  | 1.50 | <0.100 | <0.100 | <0.100 | PCB 208 |  | 1.64 |  | 1.54 | <0.100 | 0.0629 |  | 0.064 | PCB 208 |  | 1.85 | <0.100 | <0.100 | PCB 208 |  | <0.100 |  | 0.0598 | PCB 208 | 0.0886 |  | <0.100 |  | <LOQ | <LOQ | 0.0913 |  | <LOQ | [9 empty] | PCB 208 | <0.100 | 0.089 | 0.092 | 0.095 |  | 1.405 | 1.42 | 1.53 | [13 empty] | PCB 208 | 0.165 | <0.100 | 0.073 | 0.072 | 0.074
PCB 209 |  | PCB 209 | 1.22 | 1.11 | 1.06 | 1.2 | 1.2 | 1.158 | 0.069426 | 1.30 + 0.21 |  | 1.0201 | 0.012075 | PCB 209 | 1.30 + 0.21 | 1.12 | <0.100 | 0.456 | 0.409 | PCB 209 | 1.30 + 0.21 | 0.973 |  | 1 | <0.100 | 0.103 |  | 0.0878 | PCB 209 | 1.30 + 0.21 | 1.11 | <0.100 | <0.100 | PCB 209 |  | <0.100 |  | 0.166 | PCB 209 | 0.0609 |  | <0.100 |  | 0.0896 | 0.0811 | 0.0799 |  | 0.0835 | 0.0053 | 6.3 | [7 empty] | PCB 209 | <0.100 | 0.14 | 0.136 | 0.14 | 1.30 + 0.21 | 1.022 | 1.027 | 0.968 | [13 empty] | PCB 209 | 0.026 | <0.100 | 0.075 | 0.077 | 0.075
PCB 8 | [11 empty] | 0.05 | 0.05 | [6 empty] | PCB 8 |  | <LOQ |  | <LOQ | <0.100 | <LOQ |  | <LOQ | [5 empty] | PCB 8 |  | <0.100 |  | <LOD | PCB 8 | 0.00709 |  | <0.100 |  | <LOD | <LOD | <LOD |  | <LOD | [9 empty] | PCB 8 | <0.100 | 0.04 | 0.039 | 0.044 |  | 0.04 | 0.043 | 0.04
PCB 29 | [11 empty] | 0.05 | 0.05
PCB 45 | [11 empty] | 0.19555 | 0.05
PCB 50 | [11 empty] | 0.05 | 0.05
PCB 77 | [11 empty] | 0.27299 | 0.04354 | [44 empty] | PCB 77 |  | 0.006 | 0.007 | 0.006 | 0.327 + 0.025 | 0.339 | 0.327 | 0.33
PCB 79 | [11 empty] | 0.14618 | 0.05
PCB 82 | [11 empty] | 0.23455 | 0.079571
PCB 92 | [11 empty] | 7.8898 | 0.12777 | [15 empty] | PCB 92 |  | 7.72 | <0.100 | <0.100
PCB 104 | [11 empty] | 0.146 | 0.05
PCB 106 | [11 empty] | 0.37705 | 0.05
PCB 121 | [11 empty] | 0.28273 | 0.13696
PCB 126 | [11 empty] | 0.40363 | 0.05 | [44 empty] | PCB 126 |  | <LOQ | <LOQ | <LOQ | 0.380 + 0.017 | <LOQ | 0.908 | <LOQ
PCB 127 | [11 empty] | 0.05 | 0.05
PCB 169 | [11 empty] | 0.078247 | 0.018128 | [44 empty] | PCB 169 |  | 0.055 | 0.057 | 0.058 | 0.106 + 0.014 | 1.028 | 1.082 | 1.12
PCB 205 | [11 empty] | 0.92001 | 0.084459 | [6 empty] | PCB 205 |  | 0.779 |  | 0.797 | <0.100 | 0.0178 |  | <LOD | [5 empty] | PCB 205 |  | <0.100 |  | 0.0257 | PCB 205 | 0.0304 |  | <0.100 |  | 0.0458 | 0.0397 | 0.0402 |  | 0.0419 | 0.0034 | 8.1 | [7 empty] | PCB 205 | <0.100 | 0.028 | 0.026 | 0.026 |  | 0.696 | 0.666 | 0.675 | [13 empty] | PCB 205 | 0.031 | <0.100 | #VALUE! | 0.013 | 0.015
[22 empty] | 297-336 | 464-513 | 676-871 |  | 297-336 | 464-513 | 676-871 |  | SRM 1946 Certified | SRM | Murre Egg CM | Murre
[14 empty] | Compound | or Reference Values | 1946 | Reference (CHS) Values | Batch 1 | Batch 2 |  | SRM 1946 Certified | SRM 1946 |  |  | Murre Egg CM | Murre Egg CM |  |  | Compound | or Reference Values | 1946 | Reference (CHS) Values | Egg CM |  |  | 1339-1416 | 1234-1335 | [5 empty] | Murre Egg Control Material |  | Eggs 1419-1441 | [6 empty] | Eggs 1443-1486 | [7 empty] | 1577-1617 |  |  |  | 1487-1503; 1535-1541; 1609-1619 |  |  | 1506-1531; 1542-1547 |  |  |  | 1577-1617 |  |  | 1487-1503; 1535-1541; 1609-1619 |  |  | 1506-1531; 1542-1547 | [5 empty] | Eggs 114-120; 189-191; 406-418; 1010-1018; 1087-1100; 1127-1136 |  |  | Eggs 275-284; 455-462;472-476;845-853; 983-988; 997-1005; 1069-1085
[20 empty] | Compound | or Reference Values | Batch 1 | Batch 2 | Batch 3 | Reference Values | Batch 1 | Batch 2 | Batch 3 | [5 empty] | Compound | Consensus Values | Gull Batch | Murre Batch |  | Compound | Max LOD or LOQ |  |  | Consensus Values |  | Batch 1-1 | Batch 1-2 | Batch 1-3 |  | Mean | SD | RSD | Batch 2-1 | Batch 2-2 | Batch 2-3 | Mean | SD | RSD | ng/g wet wt after LOD: |  | EggCM-1 | EggCM-2 | EggCM-3 |  | EggCM-1 | EggCM-2 | EggCM-3 | EggCM-1.D | EggCM-2.D | EggCM-3.D | SRM 1946 Cert/Ref | 1946-1.D | 1946-2.D | 1946-3.D | 1946-1.D | 1946-2.D | 1946-3.D | 1946-1.D | 1946-2.D | 1946-3.D | Compound | Max LOD or LOQ | Consensus Values | Batch 1-1 | Batch 1-2 | Batch 1-3 | Batch 2-1 | Batch 2-2 | Batch 2-3
2,4'-DDD |  | 2,4'-DDD | 3.08 | 2.99 | 2.23 | 2.26 | 3.11 | 2.734 | 0.4487 | 2.20 + 0.25 |  | 2.1797 | 0.13671 | 2,4'-DDD | 2.20 + 0.25 | 1.08 | <0.100 | <0.100 | <0.100 | 2,4'-DDD | 2.20 + 0.25 | 1.18 |  | 1.16 | <0.100 | <LOD |  | <LOD | 2,4'-DDD | 2.20 + 0.25 | 0.838 | <0.100 | 0.178 | 2,4'-DDD | <0.100 | <LOQ |  |  | 2,4'-DDD | 0.0825 |  |  | <0.100 |  | <LOD | <LOD | <LOD |  | <LOD | [8 empty] | 2,4'-DDD | <0.100 | <LOD | <LOD | <LOD | [7 empty] | 2.20 + 0.25 | 2.1203 | 2.2353 | 2.1204
2,4'-DDE |  | 2,4'-DDE | 1.03 | 1.39 | 1.06 | 0.7 | 0.78 | 0.992 | 0.27142 |  | 1.04 + 0.29 | 0.99773 | 0.15778 | 2,4'-DDE | 1.04 + 0.29 | 1.03 | <0.7; 0.569 ± 0.048 | 0.134 | <0.100 | 2,4'-DDE | 1.04 + 0.29 | 1.14 |  | 1.07 | <0.7 | 0.155 |  | 0.142 | 2,4'-DDE | 1.04 + 0.29 | 1.25 | <0.7; 0.569 ± 0.048 | 0.178 | 2,4'-DDE | 0.569 ± 0.048 | <LOD | [21 empty] | 2,4'-DDE | 0.569 ± 0.048 | 0.15115 | 0.15338 | 0.16352 | [7 empty] | 1.04 + 0.29 | 1.1148 | 1.184 | 1.0781 | [6 empty] | 2,4'-DDE | 0.217 | <0.100 | 0.264 | 0.247 | 0.246
2,4'-DDT |  | 2,4'-DDT | 24.7 | 20.2 | 19.8 | 19.9 | 21.6 | 21.24 | 2.0647 |  | 22.3 + 3.2 | 21.1753 | 0.05 | 2,4'-DDT | 22.3 + 3.2 | 21.0 | <1; 0.108 ± 0.10 | <0.100 | <0.100 | 2,4'-DDT | 22.3 + 3.2 | 23 |  | 23.1 | <1; 0.108 ± 0.10 | <LOD |  | <LOD | 2,4'-DDT | 22.3 + 3.2 | 21.0 | <1; 0.108 ± 0.10 | <0.100 | 2,4'-DDT | 0.108 ± 0.10 | <LOD |  |  | 2,4'-DDT | 0.393 |  |  | 0.108 ± 0.10 |  | <LOQ | <LOQ | <LOQ |  | <LOQ | [8 empty] | 2,4'-DDT | 0.108 ± 0.10 | <LOD | <LOD | <LOD | [7 empty] | 22.3 + 3.2 | 19.8913 | 19.847 | 19.4501 | [6 empty] | 4,4,'-DDD+2,4'-DDT | 0.218 | 0.108 ± 0.10 | <LOD | <LOD | <LOD
4,4,'-DDD |  | 4,4'-DDD | 16.2 | 20.4 | 17.2 | 13.1 | 14.8 | 16.34 | 2.7437 | 17.7 + 2.8 |  | 16.3477 | 0.05 | 4,4,'-DDD | 17.7 + 2.8 | 15.3 | <3 | 0.615 | 0.780 | 4,4,'-DDD | 17.7 + 2.8 | 12.2 |  | 11.2 | <3 | <LOQ |  | <LOQ | 4,4,'-DDD | 17.7 + 2.8 | 15.1 | <3 | 0.105 | 4,4,'-DDD |  | <LOD |  |  | 4,4,'-DDD | 0.0833 |  |  |  |  | <LOD | <LOD | <LOD |  | <LOD | [8 empty] | 4,4,'-DDD |  | 0.63744 | 0.66806 | 0.66696 | [7 empty] | 17.7 + 2.8 | 16.5607 | 17.4762 | 17.8866
4,4'-DDE |  | 4,4'-DDE | 327 | 339 | 354 | 330 | 331 | 336.2 | 10.895 | 373 + 48 |  | 372.3201 | 78.7832 | 4,4'-DDE | 373 + 48 | 297 | 69.5 ± 6.3 | 64.2 | 57.0 | 4,4'-DDE | 373 + 48 | 380 |  | 378 | 69.5 ± 6.3 | 66.5 |  | 66.4 | 4,4'-DDE | 373 + 48 | 288 | 69.5 ± 6.3 | 77.5 | 4,4'-DDE | 69.5 ± 6.3 | 67.7 |  |  | 4,4'-DDE | 0.381 |  |  | 69.5 ± 6.3 |  | 66.4 | 66.4 | 66.1 |  | 66.3 | 0.16 | 0.23 | [6 empty] | 4,4'-DDE | 69.5 ± 6.3 | 63.0852 | 64.7885 | 64.1856 | [7 empty] | 373 + 48 | 341.7654 | 349.0413 | 345.5476 | [6 empty] | 4,4'-DDE | 0.987 | 69.5 ± 6.3 | 68.16 | 64.15 | 66.44
4,4'-DDT |  | 4,4'-DDT | 35.4 | 32.6 | 34.31 | 30.9 | 38.5 | 34.34 | 2.885 | 37.2 + 3.5 |  | 39.1774 | 0.05 | 4,4'-DDT | 37.2 + 3.5 | 37.1 | <1 | <0.100 | <0.100 | 4,4'-DDT | 37.2 + 3.5 | 45 |  | 46 | <1 | <LOD |  | <LOD | 4,4'-DDT | 37.2 + 3.5 | 37.3 | <1 | <0.100 | 4,4'-DDT | <0.100 | <LOQ |  |  | 4,4'-DDT | 0.0823 |  |  | <0.100 |  | <LOD | <LOD | <LOD |  | <LOD | [8 empty] | 4,4'-DDT | <0.100 | <LOD | <LOD | <LOD | [7 empty] | 37.2 + 3.5 | 39.2333 | 38.4183 | 39.2409 | [6 empty] | 4,4'-DDT | 0.215 | <0.100 | 0.097 | 0.088 | 0.095
alpha-HCH |  | a-HCH | 5.31 | 5.1 | 4.75 | 5.67 | 5.85 | 5.336 | 0.44032 | 5.72 + 0.65 |  | 5.4308 | 1.1949 | alpha-HCH | 5.72 + 0.65 | 7.32 | 1.16 ± 0.18 | 1.48 | 1.53 | a-HCH | 5.72 + 0.65 | 5.83 | 5.86 | 5.99 | 1.16 ± 0.18 | 1.11 | 1.14 | 0.985 | alpha-HCH | 5.72 + 0.65 | 5.42 | 1.16 ± 0.18 | 1.17 | alpha-HCH | 1.16 ± 0.18 | 1.2 |  |  | α-HCH | 0.0846 |  |  | 1.16 ± 0.18 |  | 1.08 | 1.09 | 1.13 |  | 1.1 | 0.028 | 2.5 | [6 empty] | alpha-HCH | 1.16 ± 0.18 | 1.0431 | 1.0204 | 1.0225 |  | 1.0981 | 1.0012 | 1.0178 | 1.1369 | 1.1046 | 1.0144 | 5.72 ± 0.65 | 5.6048 | 5.3798 | 5.2016 | 5.2105 | 5.6111 | 5.5082 | 5.7187 | 5.4351 | 5.2453 | α-HCH | 0.216 | 1.16 ± 0.18 | 1.17 | 1.198 | 1.159
beta-HCH | [11 empty] | 0.53371 | 28.0525 | beta-HCH |  | 0.283 | 24.9 ± 4.3 | 32.8 | 27.1 | b-HCH |  | 0.558 | 0.458 | 0.503 | 24.9 ± 4.3 | 21.1 | 25.2 | 23.3 | beta-HCH |  | <0.100 | 24.9 ± 4.3 | 27.2 | beta-HCH | 24.9 ± 4.3 | 24 |  |  | β-HCH | 0.0841 |  |  | 24.9 ± 4.3 |  | 24.1 | 26.9 | 26.1 |  | 25.7 | 1.4 | 5.5 | [6 empty] | beta-HCH | 24.9 ± 4.3 | 25.6917 | 24.9375 | 26.3083 |  | 25.5814 | 25.0588 | 26.2019 | 24.5919 | 25.1539 | 25.8017 |  | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | β-HCH | 0.354 | 24.9 ± 4.3 | 24.14 | 22.86 | 22.27
gamma-HCH |  | g-HCH | 1.26 | 0.762 | 0.857 | 1.48 | 0.973 | 1.0664 | 0.2974 | 1.14 + 0.18 |  | 1.2842 | 0.32459 | gamma-HCH | 1.14 + 0.18 | 1.18 | <0.4; 0.342 ± 0.024 | 0.327 | 0.312 | g-HCH | 1.14 + 0.18 | 1.1 | 1.1 | 0.929 | <0.4; 0.342 ± 0.024 | <LOQ | <LOQ | <LOQ | gamma-HCH | 1.14 + 0.18 | 1.29 | <0.4; 0.342 ± 0.024 | 0.377 | gamma-HCH | 0.342 ± 0.024 | 0.363 |  |  | γ-HCH | 0.0761 |  |  | 0.342 ± 0.024 |  | 0.338 | 0.349 | 0.336 |  | 0.341 | 0.0074 | 2.2 | [6 empty] | gamma-HCH | 0.342 ± 0.024 | 0.31027 | 0.31465 | 0.3357 |  | 0.33651 | 0.31985 | 0.33307 | 0.33129 | 0.32306 | 0.32889 | 1.14 ± 0.18 | 1.1954 | 1.2771 | 1.2968 | 1.2504 | 1.2181 | 1.1616 | 1.2313 | 1.2765 | 1.2821 | γ-HCH | 0.216 | 0.342 ± 0.024 | 0.361 | 0.365 | 0.369
cis-chlordane |  | cis-chlordane | 27.2 | 31.4 | 28.8 | 24.2 | 28.7 | 28.06 | 2.634 | 32.5 + 1.8 |  | 34.8798 | 0.19088 | cis-chlordane | 32.5 + 1.8 | 32.3 | 0.254 ± 0.024 | 0.289 | 0.229 | cis-chlordane | 32.5 + 1.8 | 33.3 | 32.7 | 32.3 | 0.254 ± 0.024 | <LOQ | <LOQ | <LOQ | cis-chlordane | 32.5 + 1.8 | 36.4 | 0.254 ± 0.024 | 0.247 | cis-chlordane | 0.254 ± 0.024 | 0.177 | 0.257 |  | cis-chlordane | 0.0883 | 0.0761 |  | 0.254 ± 0.024 |  | 0.228 | 0.251 | 0.246 |  | 0.241 | 0.012 | 5 | 0.227 | 0.251 | 0.234 | 0.238 | 0.012 | 5.2 | cis-chlordane | 0.254 ± 0.024 | 0.25419 | 0.22565 | 0.24409 |  | 0.25779 | 0.23928 | 0.24062 | 0.26467 | 0.25125 | 0.23255 |  | 1.3181 | 1.348 | 1.2741 | 1.3801 | 1.286 | 1.4777 | 1.3384 | 1.3316 | 1.3838 | cis-chlordane | 0.073 | 0.254 ± 0.024 | 0.245 | 0.253 | 0.245 | 0.239 | 0.25 | 0.255
cis nonachlor |  | cis-nonachlor | 54.2 | 72.8 | 60 | 45.1 | 53.6 | 57.14 | 10.2415 | 59.1 + 3.6 |  | 47.8344 | 2.0736 | cis nonachlor | 59.1 + 3.6 | 53.1 | 1.94 ± 0.48 | 1.53 | 1.61 | cis-nonachlor | 59.1 + 3.6 | 58.4 | 62.4 | 55.3 | 1.94 ± 0.48 | 1.52 | 1.51 | 1.63 | cis nonachlor | 59.1 + 3.6 | 75.7 | 1.94 ± 0.48 | 1.84 | cis-nonachlor | 1.94 ± 0.48 | 1.9 | 2.03 |  | cis-nonachlor | 0.121 | 0.465 |  | 1.94 ± 0.48 |  | 2.2 | 2.18 | 2.18 |  | 2.19 | 0.015 | 0.7 | 2.37 | 2.44 | 2.32 | 2.38 | 0.064 | 2.7 | cis nonachlor | 1.94 ± 0.48 | 2.2906 | 2.2664 | 2.1766 |  | 2.2756 | 2.2292 | 2.1519 | 2.2185 | 2.2385 | 2.2577 | 59.1 ± 3.6 | 57.7246 | 57.6529 | 56.1346 | 58.7899 | 55.7883 | 56.2041 | 57.3119 | 57.901 | 56.451 | cis-nonachlor | 0.071 | 1.94 ± 0.48 | 2.044 | 2.17 | 1.936 | 1.957 | 2.1 | 2.101
heptachlor epoxide |  | heptachlor epoxide | 5.96 | 6.02 | 5.61 | 4.64 | 4.05 | 5.256 | 0.87157 | 5.50 + 0.23 |  | 5.7693 | 1.8256 | heptachlor epoxide | 5.50 + 0.23 | 5.43 | 4.41 ± 0.3 | 4.35 | 4.30 | trans-chlordane | 8.36 + 0.91 | 20.5 | 20.4 | 20.4 | <2 | <LOD | <LOD | 0.0827 | heptachlor epoxide | 5.50 + 0.23 | 4.08 | 4.41 ± 0.3 | 4.54 | trans-chlordane | <0.100 | 0.145 | 0.149 |  | heptachlor epoxide | 0.516 | 1.09 |  | 4.41 ± 0.3 |  | 4.45 | 4.25 | 4.32 |  | 4.34 | 0.1 | 2.3 | 4.74 | 4.73 | 4.45 | 4.64 | 0.16 | 3.5 | heptachlor epoxide | 4.41 ± 0.3 | 4.1621 | 4.4305 | 4.2593 |  | 4.1748 | 4.4219 | 4.3489 | 4.3263 | 4.4141 | 4.2882 | 5.5 ± 0.23 | 5.8695 | 5.8778 | 5.65 | 5.7508 | 5.7615 | 5.6529 | 5.8897 | 5.7528 | 5.5982 | heptachlor epoxide | 0.715 | 4.41 ± 0.3 | 4.428 | 4.383 | 4.338 | 4.562 | 4.259 | 4.185
oxychlordane |  | oxychlordane | 18 | 20.11 | 19.4 | 19.8 | 21.5 | 19.8 | 1.3 | 18.9 + 1.5 |  | 18.3768 | 7.0807 | oxychlordane | 18.9 + 1.5 | 18.6 | 7.16 ± 0.7 | 7.07 | 6.06 | trans-nonachlor | 99.6 + 7.6 | 105 | 103 | 106 | 0.481 ± 0.1 | 0.367 | 0.378 | 0.434 | oxychlordane | 18.9 + 1.5 | 20.9 | 7.16 ± 0.7 | 7.03 | trans-nonachlor | 0.481 ± 0.1 | 0.402 | 0.4 |  | oxychlordane | 0.0807 |  |  | 7.16 ± 0.7 |  | 6.73 | 6.72 | 6.68 |  | 6.71 | 0.027 | 0.4 | [6 empty] | oxychlordane | 7.16 ± 0.7 | 6.8634 | 6.6963 | 6.7206 | [7 empty] | 18.9 + 1.5 | 17.8539 | 18.7879 | 17.4976 | [6 empty] | oxychlordane | 0.206 | 7.16 ± 0.7 | 6.46 | 7.001 | 6.793
trans-chlordane |  | trans-chlordane | 7.22 | 8.1 | 7.34 | 6.16 | 7.65 | 7.294 | 0.71936 | 8.36 + 0.91 |  | 8.4653 | 0.046379 | trans-chlordane | 8.36 + 0.91 | 8.58 | <2 | <0.100 | <0.100 | heptachlor |  | 0.146 |  | <LOD | <0.100 | <LOD |  | <LOD | trans-chlordane | 8.36 + 0.91 | 7.52 | <2 | 0.138 | Heptachlor | <0.100 | <LOQ |  |  | trans-chlordane | 0.0769 | 0.0769 |  | <0.100 |  | <LOD | <LOD | <LOD |  | <LOD |  |  | <LOD | <LOD | <LOD | <LOD |  |  | trans-chlordane | <0.100 | <LOD | <LOD | <LOD |  | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD | 8.36 ± 0.91 | 8.4313 | 8.6484 | 9.0179 | 8.4048 | 8.0749 | 8.6189 | 8.2369 | 8.0577 | 8.691 | trans-chlordane | 0.072 | <0.100 | <LOD | <LOD | <LOD | <LOD | <LOD | <LOD
trans-nonachlor |  | trans-nonachlor | 91.3 | 107 | 104 | 83.1 | 95 | 96.08 | 9.6756 | 99.6 + 7.6 |  | 93.1239 | 0.47483 | trans-nonachlor | 99.6 + 7.6 | 106 | 0.481 ± 0.1 | 0.449 | 0.302 | heptachlor epoxide | 5.50 + 0.23 | 5.83 | 5.9 | 5.74 | 4.41 ± 0.3 | 4.11 | 4.11 | 4.21 | trans-nonachlor | 99.6 + 7.6 | 112 | 0.481 ± 0.1 | 0.528 | heptachlor epoxide | 4.41 ± 0.3 | 3.47 | 3.97 |  | trans-nonachlor | 0.0869 | 0.0767 |  | 0.481 ± 0.1 |  | 0.438 | 0.442 | 0.431 |  | 0.437 | 0.0055 | 1.3 | 0.423 | 0.42 | 0.411 | 0.418 | 0.0065 | 1.6 | trans-nonachlor | 0.481 ± 0.1 | 0.45732 | 0.44356 | 0.48479 |  | 0.47643 | 0.43628 | 0.4755 | 0.46118 | 0.43428 | 0.47444 | 99.6 ± 7.6 | 98.4303 | 94.0667 | 99.1891 | 100.3872 | 100.6857 | 99.8159 | 98.8135 | 97.8897 | 100.1693 | trans-nonachlor | 0.073 | 0.481 ± 0.1 | 0.465 | 0.46 | 0.428 | 0.428 | 0.455 | 0.418
dieldrin |  | dieldrin | 29.9 | 34.4 | 33.2 | 26.8 | 24.2 | 29.7 | 4.2732 | 32.5 + 3.5 |  | 32.7101 | 3.1444 | dieldrin | 32.5 + 3.5 | 29.9 | 3.34 ± 0.36 | 3.24 | 3.00 | oxychlordane | 18.9 + 1.5 | 18.6 | 18.7 | 18.4 | 7.16 ± 0.7 | 7.76 | 7.57 | 7.38 | dieldrin | 32.5 + 3.5 | 32.5 | 3.34 ± 0.36 | 3.29 | oxychlordane | 7.16 ± 0.7 | 6.55 | [43 empty] | dieldrin | 1.484 | 3.34 ± 0.36 | 3.427 | 3.333 | 3.572
aldrin | [11 empty] | 0.43019 | 0.05 | aldrin |  | 1.45 | <0.100 | <0.100 | <0.100 | aldrin |  | <LOD |  | <LOD | <0.100 | <LOD |  | <LOD | aldrin |  | <0.100 | <0.100 | <0.100
endrin | [11 empty] | 5.2926 | 0.017007 | [6 empty] | dieldrin | 32.5 + 3.5 | 32.6 |  | 33.8 | 3.34 ± 0.36 | 3.24 |  | 3.4 | endrin |  | 0.963 | 0.452 ± 0.082 | 0.219
endosulfan II | [11 empty] | 1.9574 | 0.18191 | [6 empty] | endosulfan I |  | 30.8 | 27.3 | 27.4 | 0.271 ± 0.049 | 0.253 | 0.26 | 0.222 | endosulfan II |  | 28.4 | 0.421 ± 0.076 | 0.432 | [46 empty] | endosulfan II | 0.361 | 0.421 ± 0.076 | 0.423 | 0.42 | 0.422
endosulfan sulfate | [11 empty] | 0.040846 | 0.0066578 | endosulfan sulfate |  | 0.312 | <0.100 | <0.100 | <0.100 | endosulfan II |  | 50.1 | 50.7 | 54.5 | 0.421 ± 0.076 | 4.92 | 5.56 | 5.4 | endosulfan sulfate |  | <0.100 | <0.100 | <0.100 | [46 empty] | endosulfan sulfate | 0.071 | <0.100 | <LOD | <LOD | <LOD
HCB |  | HCB | 10.1 | 7.75 | 7.33 | 7.4 | 7.82 | 8.08 | 1.1491 | 7.25 + 0.83 |  | 8.6465 | 37.9306 | HCB | 7.25 + 0.83 | 6.95 | 34.0 ± 5.3 | 41.8 | 32.9 | endosulfan sulfate |  | 0.243 | 0.219 | 0.238 | <0.100 | <LOD | <LOD | <LOD | HCB | 7.25 + 0.83 | 6.79 | 34.0 ± 5.3 | 31.9 | HCB | 34.0 ± 5.3 | 38.4 | 38.7 |  | HCB | 0.0823 |  |  | 34.0 ± 5.3 |  | 39 | 39.5 | 39.5 |  | 39.3 | 0.26 | 0.65 | [6 empty] | HCB | 34.0 ± 5.3 | 38.8878 | 39.0459 | 38.2044 | [7 empty] | 7.25 + 0.83 | 7.604 | 7.9631 | 7.96 | [6 empty] | HCB | 0.216 | 34.0 ± 5.3 | 34.51 | 36.68 | 36.32
mirex |  | mirex | 5.18 | 5.83 | 5.11 | 6.05 | 5.37 | 5.508 | 0.41306 | 6.47 + 0.77 |  | 5.9875 | 1.5679 | mirex | 6.47 + 0.77 | 5.52 | 1.58 ± 0.16 | 1.47 | 1.53 | endrin |  | 6.42 | 6.36 | 6.4 | 0.452 ± 0.082 | <LOQ | <LOQ | <LOQ | mirex | 6.47 + 0.77 | 7.49 | 1.58 ± 0.16 | 1.80 | mirex | 1.58 ± 0.16 | 1.62 |  |  | mirex | 0.0768 |  |  | 1.58 ± 0.16 |  | 1.44 | 1.46 | 1.4 |  | 1.44 | 0.028 | 2 | 1.54 | 1.55 | 1.59 | 1.6 | 0.028 | 1.81 | mirex | 1.58 ± 0.16 | 1.5876 | 1.5132 | 1.5294 | [7 empty] | 6.47 + 0.77 | 5.3498 | 5.2226 | 5.0986 | [6 empty] | mirex | 0.218 | 1.58 ± 0.16 | 1.623 | 1.59 | 1.611
Octachlorosytrene | [11 empty] | 3.8125 | 0.96423 | octachlorosytrene |  | 4.69 | 0.913 ± 0.087 | 1.00 | 0.867 | HCB | 7.25 + 0.83 | 7.81 | 7.86 | 7.64 | 34.0 ± 5.3 | 39.3 | 39.7 | 38.5 | [5 empty] | octachlorosytrene | 0.913 ± 0.087 | 0.915 |  |  | octachlorosytrene | 0.0662 |  |  | 0.913 ± 0.087 |  | 0.963 | 0.982 | 0.968 |  | 0.971 | 0.01 | 1 | [6 empty] | Octachlorosytrene | 0.913 ± 0.087 | 0.90928 | 0.92712 | 0.91796 | [8 empty] | 4.4867 | 4.4684 | 4.1489 | [6 empty] | octachlorosytrene | 0.035 | 0.913 ± 0.087 | 0.847 | 0.891 | 0.855 | 0.439 | 0.41 | 0.417
pentachlorobenzene | [11 empty] | 2.3218 | 12.3125 | pentachlorobenzene |  | 0.116 | 1.93 ± 0.5 | 3.39 | 2.66 | mirex | 6.47 + 0.77 | 6.57 | 6.54 | 6.53 | 1.58 ± 0.16 | 1.6 | 1.53 | 1.58 | pentachlorobenzene |  | 0.288 | 1.93 ± 0.5 | 1.48 | pentachlorobenzene | 1.93 ± 0.5 | 1.83 | 2.09 |  | pentachlorobenzene | 0.0569 |  |  | 1.93 ± 0.5 |  | 2.47 | 2.42 | 2.42 |  | 2.44 | 0.031 | 1.3 | [6 empty] | pentachlorobenzene | 1.93 ± 0.5 | 1.1452 | 1.7831 | 1.8242 | [8 empty] | 0.35093 | 0.36504 | 0.35122 | [6 empty] | pentachlorobenzene | 0.146 | 1.93 ± 0.5 | 1.925 | 1.984 | 1.835 | <LOD | <LOD | <LOD
[20 empty] | octachlorosytrene |  | 4.06 |  | 4.15 | 0.913 ± 0.087 | 0.888 |  | 0.816
heptachlor | [11 empty] | 0.05 | 0.05 | [6 empty] | pentachlorobenzene |  | 0.253 | 0.291 | 0.289 | 1.93 ± 0.5 | 2.13 | 2.57 | 2.36 | [51 empty] | heptachlor | 0.341 | <0.100 | 0.24 | <LOQ | <LOQ
|  | 1 Values obtained using GC-MS | [17 empty] | BDE 17 | 0.1 | <LOD |  | <LOD |  | <LOD |  | <LOD
ng/g wet wt: | [11 empty] | 1946 | MURRE | [6 empty] | BDE 25 | <0.2 | 0.0938 |  | <LOD |  | <LOD |  | <LOD
[12 empty] | 2.93 | 2.984 | [6 empty] | BDE 28 |  | 0.834 |  | 0.888 |  | <LOQ |  | <LOQ
| Certified values | [18 empty] | BDE 30 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD | [29 empty] | ng/g wet wt after LOD: | CM consensus | EggCM-1.D | EggCM-2.D | EggCM-3.D | SRM 1946 Cert/Ref | 1946-1.D | 1946-2.D | 1946-3.D
BDE28 | 0.73 with 33 | [10 empty] | 0.59689 | <0.100 | BDE 28 | 0.742+0.027 (w/ BDE 33) | 0.899 |  | <0.100 | <0.100 | BDE 47 | 29.9+2.3 | 30.7 |  | 29.8 | 0.771 ± 0.14 | 0.777 |  | 0.795 | BDE 28 | 0.742+0.027 | 0.877 |  | 0.499 | [24 empty] | BDE28 |  | <LOD | <LOD | <LOD | 0.742+0.027 | 0.759 | 0.748 | 0.774
BDE47 | 27.8 | [10 empty] | 26.2809 | 0.70399 | BDE 47 | 29.9+2.3 | 34.0 | 0.771 ± 0.14 | 0.856 | 0.747 | BDE 49 | 1.10+0.23 | 0.942 |  | 0.967 |  | <LOQ |  | <LOQ | BDE 47 | 29.9+2.3 | 29.9 | 0.771 ± 0.14 | 0.788 | BDE47 |  | 0.771 ± 0.14 |  | 0.817 | BDE47 | 0.0874 |  | 0.771 ± 0.14 |  | 0.734 | 0.756 | 0.711 |  | 0.733 | 0.023 | 3.1 | [7 empty] | BDE47 | 0.771 ± 0.14 | 0.795 | 0.779 | 0.823 | 29.9+2.3 | 29.548 | 30.816 | 29.779 | [13 empty] | BDE 47 | 0.124 | 0.771 ± 0.14 | 0.67 | 0.678 | 0.682
BDE66 | 1.2 | [10 empty] | 1.0797 | <0.100 | BDE 66 | 1.35+0.16 | 1.40 |  | <0.100 | <0.100 | BDE 66 | 1.35+0.16 | 2 |  | 2.08 |  | <LOD |  | <LOD | BDE 66 | 1.35+0.16 | 1.41 |  | 0.432 | [24 empty] | BDE66 |  | #VALUE! | #VALUE! | #VALUE! | 1.35+0.16 | 1.831 | 1.889 | 1.819
BDE99 | 17.3 | [10 empty] | 17.777 | 0.30347 | BDE 99 | 18.5+2.1 | 17.7 | 0.530 ± 0.21 | 0.628 | 0.516 | BDE 71 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD | BDE 99 | 18.5+2.1 | 19.3 | 0.530 ± 0.21 | 0.687 | BDE99 |  | 0.530 ± 0.21 |  | <LOQ | BDE99 | 0.0797 |  | 0.530 ± 0.21 |  | 0.28 | 0.279 | 0.241 |  | 0.267 | 0.022 | 8.3 | [7 empty] | BDE99 | 0.530 ± 0.21 | 0.468 | 0.422 | 0.418 | 18.5+2.1 | 18.318 | 18.163 | 17.768 | [13 empty] | BDE 99 | 0.136 | 0.530 ± 0.21 | 0.427 | 0.425 | 0.415
BDE100 | 8 | [10 empty] | 7.3453 | 0.22218 | BDE 100 | 8.57+0.52 | 8.01 | 0.310 ± 0.13 | 0.253 | 0.261 | BDE 74 |  | 0.213 |  | 0.214 |  | <LOD |  | <LOD | BDE 100 | 8.57+0.52 | 5.77 | 0.310 ± 0.13 | 0.456 | BDE100 |  | 0.310 ± 0.13 |  | 0.479 | BDE100 | 0.0873 |  | 0.310 ± 0.13 |  | 0.303 | 0.291 | 0.27 |  | 0.288 | 0.017 | 5.8 | [7 empty] | BDE100 | 0.310 ± 0.13 | 0.286 | 0.292 | 0.32 | 8.57+0.52 | 6.661 | 6.436 | 6.26 | [13 empty] | BDE 100 | 0.23 | 0.310 ± 0.13 | 0.25 | 0.254 | 0.251
BDE138 | <0.2 | [10 empty] | <0.100 | <0.100 | BDE 138 | <0.2 | <0.100 | 0.644 ± 0.087 | 0.973 | 0.714 | BDE 75 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD | BDE 138 | <0.2 | 0.140 | 0.644 ± 0.087 | 0.682 | [24 empty] | BDE138 | 0.644 ± 0.087 | 0.716 | 0.711 | 0.726 |  | #VALUE! | #VALUE! | #VALUE!
BDE153 | 2.7 | [10 empty] | 2.841 | <0.100 | BDE 153 | 2.81+0.41 | 2.66 |  | <0.100 | <0.100 | BDE 85 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD | BDE 153 | 2.81+0.41 | 4.05 |  | 0.218 | [24 empty] | BDE153 |  | #VALUE! | #VALUE! | #VALUE! | 2.81+0.41 | 2.92 | 2.878 | 2.949 | [13 empty] | BDE 153 | 0.121 |  | <LOD | <LOD | <LOD
BDE154 | 5.5 | [10 empty] | 5.1907 | <0.100 | BDE 154 | 5.77+0.80 | 5.09 |  | <0.100 | <0.100 | BDE 97+118 |  | 0.505 |  | 0.563 |  | <LOD |  | <LOD | BDE 154 | 5.77+0.80 | 6.86 |  | 0.992 | [24 empty] | BDE154 |  | #VALUE! | #VALUE! | #VALUE! | 5.77+0.80 | 5.704 | 5.78 | 5.886 | [13 empty] | BDE 154 | 0.096 |  | <LOD | <LOD | <LOD
BDE155 | 0.5 | [10 empty] | 0.28803 | -0.063498 | [6 empty] | BDE 99 | 18.5+2.1 | 17.3 |  | 17.2 | 0.530 ± 0.21 | <LOQ |  | <LOQ | BDE 155 | 0.51+0.11 | 0.410 |  | <0.100 | [24 empty] | BDE155 |  | #VALUE! | #VALUE! | #VALUE! | 0.51+0.11 | 0.465 | 0.505 | 0.472 | [13 empty] | BDE 155 | 0.062 |  | <LOQ | <LOQ | <LOQ
BDE_183 | <0.2 | [10 empty] | <0.100 | <0.100 | [6 empty] | BDE 100 | 8.57+0.52 | 8.23 |  | 8.43 | 0.310 ± 0.13 | <LOQ |  | <LOQ | BDE 183 | 0.235+0.033 | <0.100 |  | 0.106 | [24 empty] | BDE_183 |  | #VALUE! | #VALUE! | #VALUE! | 0.235+0.033 | #VALUE! | #VALUE! | #VALUE!
[20 empty] | BDE 101 |  | 0.48 |  | 0.46 |  | <LOD |  | <LOD
BDE17 | 0.1 | [10 empty] | <0.100 | <0.100 | [6 empty] | BDE 116 | <0.2 | <LOQ |  | <LOQ |  | <LOQ |  | <LOQ | [29 empty] | BDE17 |  | #VALUE! | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE!
BDE33 | [11 empty] | #NUM! | #NUM! | [6 empty] | BDE 119 |  | 0.364 |  | 0.337 |  | <LOD |  | <LOD | BDE 33 |  | <0.100 |  | <0.100
[20 empty] | BDE 138 | <0.2 | <LOQ |  | <LOQ | 0.644 ± 0.087 | <LOQ |  | <LOQ
BDE30 | <0.2 | [10 empty] | #NUM! | #NUM! | [6 empty] | BDE 139 |  | <LOQ |  | <LOQ |  | <LOQ |  | <LOQ | [29 empty] | BDE30 |  | <LOD | #VALUE! | #VALUE! |  | #VALUE! | #VALUE! | #VALUE!
BDE25 | <0.2 | [10 empty] | -0.019318 | 0.063181 | [6 empty] | BDE 153 | 2.81+0.41 | 2.83 |  | 2.98 |  | <LOQ |  | <LOQ | BDE 49 | 1.10+0.23 | 1.20 |  | 0.682 | [24 empty] | BDE25 |  | 0.157 | 0.168 | 0.155 |  | 0.501 | 0.514 | 0.521
BDE_49 | 1 | [10 empty] | 0.55438 | -0.013638 | [6 empty] | BDE 154 | 5.77+0.80 | 5.21 |  | 5.45 |  | <LOQ |  | <LOQ | [29 empty] | BDE_49 |  | #VALUE! | #VALUE! | #VALUE! | 1.10+0.23 | 1.107 | 1.119 | 1.104
[20 empty] | BDE 155 | 0.51+0.11 | 0.517 |  | 0.494 |  | <LOQ |  | <LOQ
BDE_75 | <0.2 | [10 empty] | #NUM! | #NUM! | [6 empty] | BDE 156 | <0.2 | <LOQ |  | <LOQ |  | <LOQ |  | <LOQ
BDE85 | <0.2 | [10 empty] | <0.100 | 0.53073 | [6 empty] | BDE 173+190 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD
BDE116 | <0.2 | [10 empty] | 0.1869 | -0.079759 | [6 empty] | BDE 181 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD
BDE156 | <0.2 | [10 empty] | #NUM! | #NUM! | [6 empty] | BDE 182 |  | <LOD |  | <LOD |  | <LOD |  | <LOD
BDE_191 | <0.2 | [10 empty] | #NUM! | #NUM! | [6 empty] | BDE 183 | 0.235+0.033 | 0.235 |  | 0.22 |  | <LOQ |  | <LOQ
BDE_181 | <0.2 | [10 empty] | #NUM! | #NUM! | [6 empty] | BDE 185 |  | <LOD |  | <LOD |  | <LOD |  | <LOD
BDE71 | <0.2 | [10 empty] | <0.100 | <0.100 | [6 empty] | BDE 191 | <0.2 | <LOD |  | <LOD |  | <LOD |  | <LOD
BDE_190 | <0.2 | [10 empty] | <0.100 | <0.100 | [6 empty] | BDE 196 |  | <LOQ |  | <LOQ |  | <LOQ |  | <LOQ
[20 empty] | BDE 198 |  | <LOQ |  | <LOQ |  | <LOQ |  | <LOQ
[20 empty] | BDE 203 |  | <LOQ |  | <LOQ |  | <LOQ |  | <LOQ
[20 empty] | BDE 204+197 |  | <LOD |  | <LOD |  | <LOD |  | <LOD

## Mercury
|  |  | Value | (ng/g) | [5 empty] | 2010 LAAL - DMA 80 Measurements (Determination of Hg) The mass fraction of Hg was determined with a direct mercury analyzer DMA 80 (Milestone Scientific, Shelton, CT). External calibration utilizing SRM 1946 Lake Superior Fish Tissue was performed once prior to performing sample analysis of standards and unknown samples. The external calibration curve was created by aliquoting different masses of the fresh frozen SRM into nickel sample boats. Bracketing standards of QC04ERM1, SRM 1946, and NSKI17LAAL10 were used in between blocks of 7-8 unknown samples in order to verify the instrument calibration and monitor for drift. The samples were then dried at 300 ºC for 60 s, thermally decomposed at 850 ºC for 180 s, catalytically reduced to Hg0, and trapped on a gold amalgamation trap. The standard drying time of 60 s was increased for some samples according to the following equation: Drying time (s) = sample mass (mg) x fraction moisture in sample x 0.60 The mercury was then thermally desorbed and the Hg atomic absorption measured at 254 nm for the unknown Laysan Albatross egg samples, and the QC04ERM1 and SRM 1946 control samples. | 1999-2002 Mercury Samples were run in 12 analytical batches of six samples, each of which consisted of four eggs, one control sample and one method blank. The sample dissolution procedures used multiple iterations of microwave digestion. For each digestion batch, approximately 0.9 g sample aliquots of homogenized egg tissue were digested along with one approximately 0.6 g control sample aliquot of SRM 2976 Mussel Tissue (Trace Elements and Methylmercury) and a procedural blank. The certified value for total Hg in SRM 2976 is 0.0610 μg/g ± 0.0036 μg/g. Each analytical sample was spiked with a known quantity of isotopically enriched mercury (201Hg enriched isotopic spike) prior to addition of the 5 mL nitric acid decomposition medium. The amount of spike delivered to the blank vessels was reduced to curtail errors due to overspiking. All samples were digested in a Perkin-Elmer (Shelton, CT) Multiwave microwave oven at the highest possible temperatures (up to 300 °C) and pressures (up to 8 MPa) in order to equilibrate the spiked mercury with the natural mercury present in the samples. The resulting digests were vented and diluted with high purity water to a total volume of approximately 40 mL and nonquantitatively transferred into 60 mL polyethylene bottles. Further sample dilutions ranging from 1:3 to 1:40 were required prior to analysis. SRM 3133 Mercury Spectrometric Solution was used to calibrate the isotopic spike solution prior to its use. First, an approximately 100 ng/g 201Hg isotopic spike solution was prepared using 3% (mass fraction) HNO3 and 0.5% (mass fraction) K2Cr2O7 as the diluent to reduce loss of elemental mercury. Two approximately 100 ng/g natural mercury solutions were quantitatively prepared from SRM 3133 in 3% (mass fraction) HNO3 only. The natural and enriched Hg solutions were then mixed (by mass) to obtain four spike calibration solutions having a target 201Hg/202Hg ratio of 2:1. Thus, the mean spike concentration obtained from the four spike calibration mixes was used as the working concentration for the spike solution. The spike solution was periodically re-calibrated with freshly prepared SRM 3133 as even the K2Cr2O7 stabilizer could not indefinitely postpone the liberation of elemental Hg from the spike solution. The variability in the spike calibration (expressed as the percent relative standard deviation of the working spike concentration ± 1s) as calculated from four spike calibration mixes for n = 5 separate spike calibration experiments was 0.22% ± 0.03%. ID-CV-ICPMS Measurements The mercury reduction chemistry and sample introduction system has been described previously (Christopher et al. 2001) and is only briefly summarized here. The mercury in each sample solution was reduced to elemental Hg using a reductant solution of 10% (mass fraction) SnCl2 in 7% (mass fraction) HCl in water. A gas-liquid separator was used to strip the Hg from the sample digest solutions using a stream of Ar gas (approximately 250 mL/min.). Delivery of Hg0 to the mass spectrometer was achieved by plumbing the gas output of the gas-liquid separator into the inductively coupled plasma mass spectrometer (ICPMS) injector line. A mass flow controller (AALBORG Model GFC 171, Greenwich, CT), controlled with LabView software and National Instruments (Austin, TX) data acquisition hardware, regulated gas flow through the gas-liquid separator. Numerous 201Hg/202Hg isotope ratio pairs were collected for all calibration and analytical samples using a Thermo Elemental PQ3 ICPMS operating in the time resolved analysis mode. The ICP power was maintained at 1350 W forward power and Ar gas flow rates were 13.5 L/min, 0.85 L/min and 0.81 L/min for the coolant, auxiliary and injector lines, respectively. Each analytical run produced a temporal profile that consisted of a 240 s data collection window. Measurements during the first 20 s to 40 s of the profile when a 5% (mass fraction) HNO3 wash solution (and not a sample) was present in the gas-liquid separator established the reference baselines. A typical isotope-time profile is depicted in Fig. 2 for a spiked egg sample from East Amatuli Island. The mean of eight baseline-corrected 10 s integration windows was used to establish the measured 201Hg/202Hg isotope ratio for all spike calibration, blank, control and unknown samples. All data were corrected for detector dead time and mass discrimination using the methods outlined by Vanhaecke and coworkers (1998). | 2002-2005 Isotope dilution cold vapor inductively coupled plasma mass spectrometry (ID-CV-ICPMS) was used to determine total mercury (THg) levels in the egg samples. This technique has been described in detail elsewhere (see Christopher et al. 2002), and is only summarized here. Isotopically enriched 201Hg spike solution (Oak Ridge National Laboratory) was prepared and calibrated using NIST SRM 3133 (Mercury Spectrometric Solution). The spike was then added quantitatively to a mass of sample (0.2 g to 1.0 g) to yield an isotopic ratio (201Hg/202Hg) that minimized random error propagation. Samples were then digested and equilibrated using microwave-assisted acid decomposition with high-purity nitric acid (Fisher Scientific, Suwanee, Georgia). Eggs collected before 2002 were digested in a Perkin-Elmer (Shelton, Connecticut) Multiwave microwave oven at the highest possible temperatures (up to 300 °C) and pressures (up to 8 MPa) using quartz microwave decomposition vessels and high-purity nitric acid. Egg samples collected in 2002 and in subsequent years were digested in a CEM Discover open-focus microwave programmed for a 15 watt ramp for 2 minutes, followed by a 70 watt ramp for 2 minutes, and a 70 watt hold for 8 minutes, resulting in typical digestion conditions of 100 °C and 15.5 kPa. The digestant was mixed with a SnCl2 and HCl reductant solution in a gas-liquid separator, allowing cold vapor transfer of the resulting Hg0 in a stream of argon to the inductively coupled plasma mass spectrometer (ICPMS) injector. A VG Elemental Plasma Quad 3 ICPMS (Windsford, Cheshire, United Kingdom) using typical power and gas flow was used in time resolved analysis mode for measurement of isotope ratios. | 2008-2009 Total Hg Levels We employed isotope dilution cold vapor inductively coupled plasma mass spectrometry (IDCV- ICPMS) to determine total mercury (THg) levels in the egg samples. This technique has been described elsewhere (Christopher et al. 2001) and is only summarized here. Isotopically enriched 201Hg spike solution (Oak Ridge National Laboratory, USA) was prepared and calibrated using NIST SRM 1641d Mercury in Water standard solution. The spike was added quantitatively to the sample masses to yield an isotopic ratio (201Hg/202Hg) that minimized random error propagation. The specific sequences and methods for digestion and addition of the 201Hg spike solution to the 2008 and 2009 batches are described below. The resulting digestant was mixed with a SnCl2 and HCl reductant solution in a gas-liquid separator that allowed cold vapor transfer of the resulting Hg0 in a stream of argon gas to the inductively 10 coupled plasma mass spectrometer (ICPMS) injector. We used a VG Elemental Plasma Quad 3 ICPMS (Windsford, Cheshire, United Kingdom) set at normal power and gas flow rates in its time resolved analysis mode to measure the isotope ratios. 2008 Collections, Batches 1-8: The aliquots that we digested for THg weighed about 0.5 g each and we weighed them directly into quartz microwave vessels using a three-place analytical balance (Mettler, Toledo OH, USA, Model PG603-S). The vessels contained 1.0 mL nitric acid (Optima, Fisher Scientific, Suwanee GA, USA), 0.5 mL hydrogen peroxide (Trace Select, Fluka, Germany), and 0.5 mL of high purity water. We digested the samples in a Perkin-Elmer (Shelton CT, USA) Multiwave microwave oven at the highest possible temperatures (up to 300 ºC) and pressures (up to 8 MPa) that was programmed as follows: (1) 100W to 600 W ramp up over 5 min, (2) 5 min hold at 600 W, and (3) 20 min hold at 1000 W. After digestion, we transferred the samples to clean containers, diluted them to 10.0 mL with high purity water, and blended them with aliquots of isotopically enriched mercury (201Hg enriched isotopic spike, Oak Ridge National Laboratory, USA) that had been calibrated against NIST SRM 1641d Mercury in Water standard solution. As a final step before running the samples, we diluted them to about 0.5 ng/g each with high purity water and 2 % mass fraction HNO3 and 2 % mass fraction HCl. We ran batches 1-8 during 24 February 2010 – 5 March 2010. Each batch consisted of two rotors with a total of twelve vessels containing 10 unknown egg samples, one digest of QC04- ERM1 Egg Reference Material 1, and one method blank. QC04-ERM1 is an in-house control material that was made from pooled murre eggs (Vander Pol et al. 2007). It has been analyzed repeatedly for Hg and has been assigned a reference value. We ran the QC04-ERM1 digest at the beginning and end of each batch for method validation (means are reported). We also ran a mass bias correction sample at the beginning and end of each batch to capture instrumental drift in mass bias and correct the measured Hg isotopic ratios, as appropriate. Each experiment consisted of two batches run consecutively on the ICPMS. 2009 Collections, Batches 9-10: The aliquots that we digested for THg weighed between about 0.15 g and 0.3 g each. We weighed them directly into 35 mL glass microwave vessels using a five-place analytical balance (Sartorius, Model MC210S), and then spiked them with known quantities of isotopically enriched Hg (201Hg enriched isotopic spike, Oak Ridge National Laboratory, USA) that had been calibrated against NIST SRM 1641d Mercury in Water standard solution. We digested the samples with 1.2 mL nitric acid (Optima, Fisher Scientific, Suwanee GA USA) and 0.3 mL hydrogen peroxide (Trace Select, Fluka, Germany) in a CEM (Matthews NC, USA) Discover open-focus microwave oven programmed at 50 W for 2 min, followed by a 50 W ramp up for 2 min, followed by a 225 W hold for 6 min (“Dolphin skin 35” method on the microwave laptop). Maximum temperatures and pressures were 250 ºC and 275 psi, respectively, with typical maximum digestion conditions of 200 ºC and 200 psi. After digestion, we transferred the samples to clean containers and diluted them to about 0.5 ng/g each with high purity water and 2 % mass fraction HNO3 and 2 % mass fraction HCl in preparation for running them. We ran batches 9 to 10 during 16-22 September 2010. Each batch consisted of 20 unknown egg samples, 3 digests of QC04-ERM1 Egg Reference Material 1, and two method blanks. We ran 11 3 digests of QC04-ERM1 and 3 mass bias correction samples at the beginning, middle, and end of each batch in order to validate the method and capture instrumental drift in mass bias and correct the measured isotopic ratios, as appropriate.
|  |  | Hg |  | Expanded | DF | Coverage
Eggs | Batch # | Sample | Measured | ± | Uncertainty |  | Factor | Blank | (ng)
1999-2000
2001 | 1 | SRM 2976 | 59 |  |  |  |  | 0.53
Eggs 68-72; 78-82
| 2 | SRM 2976-2 | 57.6 |  |  |  |  | 0.33
Eggs 73-77; 83-87
|  | Certified for SRM 2976 | 61 | ± | 3.6
2002-2005
|  | SRM 1947-1 | 260.9 | ± | 15.3 | 5.8 | 2.57 | 0.8
|  | SRM 1947-2 | 258.6 | ± | 15.2 | 5.8 | 2.57 | 0.8
Batch 1 |  | SRM 1947-3 | 256.2 | ± | 15.1 | 5.8 | 2.57 | 0.8
|  | CCQM-1 | 100.1 | ± | 3.5 | 18.7 | 2.1 | 1.9
|  | CCQM-2 | 99.1 | ± | 3.5 | 19 | 2.1 | 1.9
|  | CCQM-3 | 98 | ± | 3.5 | 19.4 | 2.09 | 1.9
|  | SRM 1947-1 | 256.2 | ± | 14.5 | 6.1 | 2.45 | 1
|  | SRM 1947-2 | 253.9 | ± | 14.4 | 6.1 | 2.45 | 1
|  | SRM 1947-3 | 253.3 | ± | 23.9 | 10.6 | 2.23 | 1
Batch 2 |  | CCQM-1 | 100.8 | ± | 4.2 | 43.1 | 2.02 | 2.8
|  | CCQM-2 | 99.8 | ± | 4.2 | 44.2 | 2.02 | 2.8
|  | CCQM-3 | 99.4 | ± | 9.7 | 13.6 | 2.16 | 2.9
|  | SRM 1947-1 | 299.4 | ± | 18.2 | 13.2 | 2.16 | 3
Batch 3 |  | SRM 1947-2 | 294.7 | ± | 18 | 13.5 | 2.16 | 3
|  | CCQM-1 | 98.9 | ± | 10.7 | 2206.5 | 1.96 | 9.3
|  | CCQM-2 | 98.8 | ± | 10.7 | 2213 | 1.96 | 9.3
|  | SRM 1947-1 | 255 | ± | 14.4 | 6.1 | 2.45 | 1
|  | SRM 1947-2 | 254.1 | ± | 14.4 | 6.1 | 2.45 | 1.1
Batch 4 |  | CCQM-1 | 92.8 | ± | 3.9 | 46.2 | 2.01 | 2.9
|  | CCQM-2 | 92.2 | ± | 3.9 | 47 | 2.01 | 2.9
|  | SRM 1947-1 | 260.7 | ± | 15.2 | 5.7 | 2.57 | 0.02
|  | SRM 1947-2 | 255.5 | ± | 14.9 | 5.7 | 2.57 | 0.02
|  | SRM 1947-3 | 255.3 | ± | 14.9 | 5.7 | 2.57 | 0.02
Batch 5 |  | CCQM-1 | 99.7 | ± | 3.3 | 14.1 | 2.14 | 0.1
|  | CCQM-2 | 97.8 | ± | 3.3 | 14.5 | 2.14 | 0.1
|  | CCQM-3 | 97.5 | ± | 3 | 6.2 | 2.45 | 0.1
|  | SRM 1947-1 | 256.1 | ± | 14.8 | 5.5 | 2.57 | 0.5
|  | SRM 1947-2 | 255.4 | ± | 14.8 | 5.5 | 2.57 | 0.5
Batch 6 |  | SRM 1947-3 | 256.9 | ± | 14.9 | 5.5 | 2.57 | 0.5
|  | CCQM-1 | 101.1 | ± | 3.3 | 10 | 2.26 | 1.2
|  | CCQM-2 | 100.7 | ± | 3.2 | 10 | 2.23 | 1.2
|  | CCQM-3 | 101 | ± | 3.3 | 10 | 2.26 | 1.2
|  | Certified Value ± Expanded Uncertainty | Mean Measured Value ± SD
| SRM 1947 | 253.5 ± 5.5 | 261.4 | ± | 5.5
| EGG CCQM | 100.9 ± 3.3 | 98.4 | ± | 2.1
2008 Eggs 1234-1416 | 1 | QC04-ERM1 | 100 | ± | 2.1 | 51 | 2 | 0.11
| 2 | QC04-ERM1 | 100 | ± | 2.1 | 51 | 2 | 0.113
| 3 | QC04-ERM1 | 95 | ± | 2 | 51 | 2 | 0.118
| 4 | QC04-ERM1 | 96 | ± | 2 | 51 | 2 | 0.107
| 5 | QC04-ERM1 | 101 | ± | 2.1 | 51 | 2 | 0.084
| 6 | QC04-ERM1 | 106 | ± | 2.2 | 51 | 2 | 0.081
| 7 | QC04-ERM1 | 97 | ± | 2.1 | 51 | 2 | 0.085
| 8 | QC04-ERM1 | 99 | ± | 2.1 | 51 | 2 | 0.11
2009 Eggs 1419-1486 | 9 | QC04-ERM1 | 95 | ± | 2.2 | 35 | 2 | 4.03
| 9 | QC04-ERM1 | 97 | ± | 2.2 | 38 | 2
| 9 | QC04-ERM1 | 97 | ± | 2.2 | 37 | 2 | 3.93
| 10 | QC04-ERM1 | 97 | ± | 2.1 | 52 | 2 | 0.21
| 10 | QC04-ERM1 | 99 | ± | 2.1 | 52 | 2
| 10 | QC04-ERM1 | 97 | ± | 2 | 52 | 2 | 0.22
2010 LAAL
QC04ERM1 [Hg] μg/kg
| 1 |  | 102
| 2 |  | 103
| 3 |  | 105
| 4 |  | 104
| 5 |  | 101
| 6 |  | 104
| 7 |  | 104
| 8 |  | 103
| 9 |  | 100
| 10 |  | 104
|  |  | 103 | ± | 2.7 | 72 | 2 | 0.11
| Sample | ID | Hg
|  | SRM 1946 | 439 | ± | 15
Certified
| Value |  | 433 | ± | 9
Acceptable
| Range |  | 424 | - | 442
2010 Murre
| 1 | QC04ERM1 | 102
| 2 | QC04ERM1 | 103
| 3 | QC04ERM1 | 98
| 4 | QC04ERM1 | 99
| 5 | QC04ERM1 | 98
| 6 | QC04ERM1 | 99
Determined Value |  |  | 100 | ± | 2.2 | 6.1 | 2.5 | 0.01
2011 LAAL
| 1 | QC04ERM1 | 96.3
| 2 | QC04ERM1 | 98.3
| 3 | QC04ERM1 | 95.5
| 4 | QC04ERM1 | 95.1
| 5 | QC04ERM1 | 96.9
| 6 | QC04ERM1 | 96.7
| 7 | QC04ERM1 | 96.6
| 8 | QC04ERM1 | 97.6
Determined | Value |  | 96.6 | ± | 1.4 | 104.6 | 2
2012 AK
| 1 | QC04ERM1 | 96.1
| 2 | QC04ERM1 | 98.7
| 3 | QC04ERM1 | 98.5
| 4 | QC04ERM1 | 95
| 5 | QC04ERM1 | 98.2
| 6 | QC04ERM1 | 100.1
| 7 | QC04ERM1 | 100.6
| 8 | QC04ERM1 | 99.1
| 9 | QC04ERM1 | 100.2
| 10 | QC04ERM1 | 98.4
| 11 | QC04ERM1 | 103.7
Determined | Value |  | 99 | ± | 1.6 | 13 | 2.2
2012-2014 USPI
| 1 | QC04ERM1 | 98.2
| 2 | QC04ERM1 | 94.1
| 3 | QC04ERM1 | 95.7
| 4 | QC04ERM1 | 97.5
| 5 | QC04ERM1 | 96.3
| 6 | QC04ERM1 | 97
| 7 | QC04ERM1 | 95.1
| 8 | QC04ERM1 | 97.4
| 9 | QC04ERM1 | 97.4
| 10 | QC04ERM1 | 98.2
| 11 | QC04ERM1 | 100.4
| 12 | QC04ERM1 | 95.2
| 13 | QC04ERM1 | 94.3
| 14 | QC04ERM1 | 95.1
| 15 | QC04ERM1 | 94.5
| 16 | QC04ERM1 | 99.1
| 17 | QC04ERM1 | 100.6
| 18 | QC04ERM1 | 98
| 19 | QC04ERM1 | 101.9
| 20 | QC04ERM1 | 99.8
| 21 | QC04ERM1 | 98.8
| 22 | QC04ERM1 | 97.3
| 23 | QC04ERM1 | 96.3
| 24 | QC04ERM1 | 98.6
Determined | Value |  | 97.4 | ± | 2.1 | 43 | 3 | o.4
2014 Phoebastria spp
| 1 | QC04ERM1 | 95.2
| 2 | QC04ERM1 | 95.1
| 3 | QC04ERM1 | 97.5
| 4 | QC04ERM1 | 97.2
| 5 | QC04ERM1 | 95.9
| 6 | QC04ERM1 | 97.8
| 7 | QC04ERM1 | 97.7
| 8 | QC04ERM1 | 96.5
Determined | Value |  | 96.6 | ± | 1 | 14.4 | 2.1 | 0.02
2015 BFAL
| 1 | QC04ERM1 | 100.3
| 2 | QC04ERM1 | 98.8
| 3 | QC04ERM1 | 102.5
| 4 | QC04ERM1 | 98.3
| 5 | QC04ERM1 | 97.9
| 6 | QC04ERM1 | 99.3
| 7 | QC04ERM1 | 104.4
| 8 | QC04ERM1 | 102.1
Determined | Value |  | 100.4 | ± | 2 | 8.5 | 2.3 | 0.02

## Trace Elements
2012-14 eggs
SRM | 1577c | [As] | [Cd] | [Cr] | [Pb] | [Mn] | [Se] | [Zn]
1 | 27.88 | 95.4 | -- | 63.14 | 10272.51 | 2013.83 | 181792.95
2 | 25.49 | 93.86 | -- | 61.08 | 10176.85 | 1985.98 | 179718.29
3 | 27 | 95.19 | 59.22 | 62.42 | 10176.6 | 1988.89 | 179977.09
4 | 25.24 | 97.02 | 57.43 | 62.17 | 10252.52 | 1992.99 | 181876.18
5 | 26.88 | 95.64 | 51.1 | 62.08 | 10254.71 | 1992.93 | 180458.48
6 | 25.99 | 95.17 | 51.2 | 61.86 | 10241.03 | 1976.34 | 181528.89
Determined | Value | 26.41 | 95.38 | 54.74 | 62.13 | 10229.04 | 1991.83 | 180891.98
Standard | Deviation | 1.01 | 1.01 | 4.21 | 0.67 | 41.76 | 12.41 | 957.97
% | RSD | 3.83 | 1.06 | 7.68 | 1.08 | 0.41 | 0.62 | 0.53
Uncertainty | Budget
Type | A | Absolute | (μg/kg)
Replication | 0.41 | 0.41 | 2.1 | 0.28 | 17.05 | 5.06 | 391.09
Blank | Correction | 1.51 | 0.04 | 0.84 | 0.02 | 1.26 | 3.72 | 23.67
Type | B | Absolute | (μg/kg)
Weighing | 0.03 | 0.1 | 0.05 | 0.06 | 10.23 | 1.99 | 180.89
SRM | 3100 | Series | 0.02 | 0.13 | 0.06 | 0.04 | 8.67 | 5.01 | 184.17
Combined | Uncertainty | 1.57 | 0.45 | 2.27 | 0.29 | 21.72 | 8.28 | 469.2 |  | 10.36
Effective | Degrees | of | Freedom | 5.75 | 6.76 | 3.94 | 5.86 | 13.19 | 27.63
Coverage | Factor | 2.57 | 2.45 | 3.18 | 2.57 | 2.16 | 2.05 | 2.23
Expanded | Uncertainty | 4.03 | 1.09 | 7.21 | 0.74 | 46.93 | 16.98 | 1045.45
SRM | 1566b | [Al] | [Cu] | [Pb] | [Mn] | [Rb] | [Se] | [Sr] | [Sn]
1 | 126705.67 | 77846.66 | 325.68 | 20029.33 | 3432.04 | 2246.98 | 6854.03 | 41.72
2 | 115388.84 | 74221.11 | 306.34 | 18997.42 | 3265.95 | 2154.97 | 6479.83 | 40.02
3 | 131633.96 | 71953.94 | 306.46 | 18623.78 | 3271.56 | 2041.3 | 6540.68 | 36.06
4 | 120431.44 | 73213.09 | 305.54 | 18834.76 | 3225.9 | 2124.53 | 6416.44 | 36.91
5 | 128152.6 | 72401.1 | 328.99 | 18535.38 | 3179.97 | 2067.88 | 6375.58 | 42.83
6 | 106845.79 | 72817.13 | 308.26 | 18591.76 | 3191.46 | 2089.97 | 6359.44 | 36.33
Determined | Value | 121526.38 | 73742.17 | 313.55 | 18935.41 | 3261.15 | 2120.94 | 6504.33 | 38.98
Standard | Deviation | 9244.28 | 2154.44 | 10.77 | 563.18 | 91.68 | 73.72 | 184.12 | 2.94
% | RSD | 7.61 | 2.92 | 3.44 | 2.97 | 2.81 | 3.48 | 2.83 | 7.55
Uncertainty | Budget
Type | A | Absolute | (μg/kg)
Replication | 3773.96 | 879.55 | 4.4 | 229.92 | 37.43 | 30.1 | 75.17 | 1.2
Blank | Correction | 9.03 | 0.88 | 0.02 | 1.26 | 0.15 | 3.72 | 0.15 | 0.9
Type | B | Absolute | (μg/kg)
Weighing | 121.53 | 73.74 | 0.31 | 18.94 | 3.26 | 2.12 | 6.5 | 0.04 |  | 0.04
SRM | 3100 | Series | 97.44 | 90.16 | 0.22 | 16.05 | 9.74 | 5.33 | 9.52
Combined | Uncertainty | 3777.18 | 887.23 | 4.41 | 231.26 | 38.81 | 30.86 | 76.05 | 1.5 | 5.24 | 9.3
Effective | Degrees | of | Freedom | 5.02 | 5.18 | 5.08 | 5.12 | 5.78 | 5.53
Coverage | Factor | 2.57 | 2.57 | 2.57 | 2.57 | 2.57 | 2.57 | 2.57 | 2.26
Expanded | Uncertainty | 9709.56 | 2280.69 | 11.35 | 594.47 | 99.77 | 79.34 | 195.49 | 3.4
Eggs 10-19 - St. Lazaria COMU 1999
2008 Eggs 1234-1416

## Hg Isotopes
Table | 5 | Quality | control | sample | mercury | isotope | values | compared | to | previously | reported | measurements | (all | values | are | reported
relative | to | NIST | SRM | 3133 | Mercury | Standard | Solution). | Data | were | obtained | by | running | the | samples | concurrently | with | unknowns
and | reference | materials | and | reported | as | mean | ± | 2 | SD. | Four | digests | of | in-house | QC-04ERM01 | Murre | Egg | Control | Material | and | 5
digests | of | NIST | SRM | 1947 | Lake | Michigan | Fish | Tissue | were | prepared. | Samples | sizes | reflect | analytical | repeats | prepared | from
digests | and | stock | solutions. | Additional | quality | control | samples | were | run | both | before | and | after | the | batches | containing | the | murre | eggs
were | run | to | validate | the | methods | (7 | aliquots | of | QC-04ERM01, | 5 | aliquots | of | SRM | 1947, | and | 15 | aliquots | of | UM-Almaden).
Reference
Material
Hg
(ng/g) | N | δ199Hg
(‰)
δ200Hg
(‰)
δ201Hg
(‰)
δ202H
g
(‰)
δ204Hg
(‰)
Δ199H
g
(‰)
Δ201H
g
(‰)
Intercomparison | Hg | Standard
(UM-Almaden)
Blum | and
Bergquist, | 2007 | 22-25 | -0.14
±0.06
-0.27
±0.04
-0.44
±0.07
-0.54
±0.08
-0.83
±0.11
-0.01
±0.02
-0.04
±0.04
Measured | 28 | -0.13
±0.13
-0.26
±0.15
-0.42
±0.17
-0.52
±0.20
-0.78
±0.29
0
±0.12
-0.03
±0.09
SRM | 1947 | Lake
Michigan | Fish
Tissue
254 | ± | 5
Point | et | al. | 2011 | 9 | 5.23
±0.29
0.6
±0.19
4.65
±0.44
0.99
±0.28 | N/A | 4.97
±0.25
3.89
±0.29
Measured | 11 | 5.61
±0.15
0.73
±0.13
5.05
±0.19
1.22
±0.21
1.73
±0.38
5.3
±0.13
4.13
±0.08
QC-04ERM01Murre
Egg | Control
Material
101 | ± | 3
Point | et | al. | 2011 | 7 | 1.5
±0.17
0.55
±0.24
1.78
±0.18
0.97
±0.26 | N/A | 1.25
±0.11
1.05
±0.10
Measured | 13 | 1.59
±0.12
0.6
±0.14
1.92
±0.16
1.09
±0.24
1.52
±0.44
1.31
±0.11
1.1
±0.10
ng/g | each
analytical
batch | was | run
0.25 | 2 | 1.57
±0.03
0.63
±0.07
1.94
±0.20
1.25
±0.15
1.73
±0.63
1.25
±0.07
1
±0.08
0.5 | 1 | 1.54 | 0.6 | 2.01 | 1.17 | 1.72 | 1.25 | 1.13
1 | 6 | 1.57
±0.13
0.55
±0.13
1.86
±0.12
0.99
±0.15
1.35
±0.35
1.31
±0.10
1.11
±0.04
1.25 | 2 | 1.64
±0.01
0.69
±0.01
2.03
±0.02
1.2
±0.01
1.66
±0.09
1.34
±0.01
1.12
±0.01
1.5 | 2 | 1.64
±0.10
0.62
±0.04
1.95
±0.03
1.09
±0.02
1.53
±0.13
1.36
±0.11
1.13
±0.04

## PFCs
Compound | SRM | 1947-1 | SRM | 1947-2 | SRM | 1947-3 | CoA | SRM | 1947 | Standards All abbreviations and sources of compounds can be found in ROA 839.02-09-302. An internal standard (IS) solution was prepared on October 29, 2008 by combining solutions of the following compounds into methanol and transferring the mixture into eight amber glass autosampler vial aliquots: 13C8-PFOA, 13C9-PFNA, 13C9-PFDA, 13C2-PFDoA, 18O2-PFBS, 13C4- PFOS, and 18O2-PFOSA. One of the aliquots prepared on October 29, 2008 was modified on January 7, 2009 by gravimetrically adding 13C2-PFUnA and modified a second time on February 19, 2009 with the addition of 13C4-PFBA, 13C2-PFHxA, and 18O2-PFHxS. Approximately 0.15 mL of the new IS aliquot was diluted on March 30, 2009 with 40 mL of methanol to achieve an approximate concentration of 15 ng/mL of 13C4-PFBA, 13C2-PFHxA, 13C8-PFOA, 13C9-PFNA, 13C9-PFDA, 13C2-PFUnA, 13C2-PFDoA, 18O2-PFBS, 18O2-PFHxS, 13C4-PFOS, and 18O2-PFOSA. The calibration series made in January 2009 from three independent weighings of neat PFCs in methanol was used. A solution of 13C2-PFOA prepared in methanol on December 31, 2008 at approximately 500 ng/g was further diluted on May 4, 2009 to approximately 15 ng/g and used as the recovery standard. Extraction Samples were extracted on July 30, 2009 using the potassium hydroxide (KOH in methanol) method as described in ROA 839.02-08-332. All labware (except plastic pipette tips) was prerinsed three times with methanol to remove PFCs before use. All additions of standards were performed gravimetrically to the nearest 0.00001 g. SRM 1947 Lake Michigan Fish Tissue, was used as the control material since it had been measured previously in ROA 839.02-09-305 and was expected to be the closest matrix match for seabird eggs. Previously reported mass fractions of PFCs in SRM 1947 are shown in Table 1. Three replicates of SRM 1947, six replicates of the murre egg CM and STAMP egg numbers 80 (Location, Species, Collection Year: St. Lazaria I., thick-billed murre [Uria lomvia], 2001), 327 (East Amatuli I., common murre [U. aalge], 2003), 611 (Noatak River Delta, glaucous gull [Larus hyperboreus], 2005), 642 (Sitka, glaucouswinged gull [L. glaucescens], 2005), 1305 (Sledge I., common murre, 2008), and 1377 (Cape Darby, glaucous gull, 2008) were extracted along with three procedural blanks and a six point calibration curve that ranged from approximately 0.0268 ng to 108 ng of each PFC. Approximately 0.5 g of the SRM, murre egg CM and egg samples were transferred to a 50 mL centrifuge tube using a stainless steel spatula. MilliQ water (0.5 mL) was added to centrifuge tubes and the blank tubes. Methanolic IS (0.2 mL) was added to each sample, which resulted in approximately 1 ng to 3 ng of each mass-labeled PFCs. Tubes were vortexed for 10 sec and then centrifuged at 2500 rpm for 5 mins. Three mL of 0.01 mol/L of KOH in methanol was added to the samples and then vortexed for 10 sec. Samples were then sonicated for 30 min. Next, samples were centrifuged at 2500 rpm for 5 min and the supernatant was transferred to a culture tube with a glass pipet. The original sample received 3 mL more of 0.01 mol/L of KOH in methanol and the extraction procedure was repeated. Extracts were evaporated with nitrogen in a TurboVap LV (Caliper Life Sciences, Hopkinton, MA; NIST property # 630262) at 7 psi and 35 oC until sample volumes were approximately 2.0 mL. For SRM 1947 aliquots, this combined evaporated supernatant was filtered using a Whatman UniPrep 0.2 μm filter (Stanford, ME) and the extracts were transferred to clean culture tubes. However, when this filtration was attempted for the murre egg CM 1 and 2 and STAMP egg 80, the extract was unable to pass through the filter. Additional methanol was added to murre egg CM 2 to attempt to get the extract to pass through the filter without success. These samples remained in the filter overnight while the remaining egg samples remained in the culture tubes. All samples were stored at -20 oC overnight. The following morning, the murre egg CM 1 and 2 and STAMP egg 80 samples were removed from the freezer and again attempted to pass through the filter without success. These samples were transferred back to the culture tubes and the remaining eggs had methanol added to create approximately 5 mL total sample volume. ENVI-Carb Clean-up Samples were attempted to be loaded onto Supelco Superclean ENVI-Carb SPE cartridges (3 mL, 250mg 120 – 400 mesh; Bellefonte, PA) using the RapidTrace workstations (US Dept of Commerce property #s CD0001057862 and CD0001460830) and a method called ENVCRB (see ROA 839.02-09-305 for details). The blanks were run, but the murre egg CM-1 resulted in a blocked syringe error after approximately 2 mL of methanol was added. The ENVI-Carb cartridge was placed on a vacuum manifold and an additional 3 mL of methanol was added to the collection tube. The remaining samples were centrifuged at 1200 rpm for 5 min to pull-down the “proteins”. The ENVI-Carb cartridges were loaded onto the vacuum manifold and filled two times (approximately 6 mL) with methanol and then two times with water. This was collected in clean culture tubes, but as the tubes were nearly full, it was disposed of before adding the samples for elution. The supernatant of the sample was added to the ENVI-Carb avoiding the “protein.” A rinse of approximately 2 mL of methanol was added to the culture tube, the tubes were centrifuged at 1200 rpm for 1 min, the supernatant added to the ENVI-Carb and this was repeated once more. An additional 1.5 fills (approximately 9 mL) of methanol was added to the ENVI-carb. A few samples (murre egg CM-2 and -6 and STAMP eggs 611 and 1305) took longer than the others to pull through the ENVI-Carb. These samples were bubbly and cloudy after evaporation (perhaps some saponification) and appeared to have some ENVI-Carb breakthrough. The other SPE fractions were evaporated with nitrogen in a TurboVap LV at 7 psi and 35 oC until samples were approximately 2 mL to 5 mL and stored at -20 oC overnight. The next day all samples were evaporated to approximately 0.5 mL. They were spiked with 0.2 mL of recovery standard, vortexed, and transferred to ASVs using glass pipets (cals, blanks, and SRM 1947 replicates). As the eggs were not successfully filtered before, all the murre egg CM replicates and STAMP eggs were transferred to Whatman filters as described above for SRM 1947 replicates. The filtered egg extracts were then quantitatively transferred to ASVs. LC/MS/MS and Quantification Methods Amounts of each PFC in the unknown samples were calculated using the linear equation of the calibration curve. Compounds were quantified using a relative response ratio to an internal standard compound that most closely matched the compound (see ROA 839.02-09-305). The limits of detection (LODs) were determined as the maximum of the lowest observable calibration solution divided by the sample mass or the average blank value plus 3 times the standard deviation which was then divided by the sample mass. Final mass fractions do not include the salt content of the neat materials; and mass fractions are totals including branched and linear isomers. Percent recoveries of internal standards were calculated by using the response and amount ratios in samples and the averaged response and amount ratios from the external standards, containing only the internal standard mixture and the recovery standard.
PFBA | <0.545 | <0.600 | <0.605
PFPeA | <2.46 | <2.71 | <2.73
PFHxA | <0.219 | <0.241 | <0.243
PFHpA | <14.4 | <15.9 | <16.0
PFOA | <0.739 | <0.814 | <0.821
PFNA | <0.850 | <0.936 | <0.943 |  |  |  | 0.2
PFDA | <0.784 | <0.863 | <0.870 |  |  |  | 0.26
PFUnA | <0.847 | <0.933 | <0.940 |  |  |  | 0.28
PFDoA | <0.803 | <0.884 | <0.891 |  |  |  | 0.2
PFTriA | <0.848 | <0.934 | <0.942
PFTA | <0.782 | <0.861 | <0.869
PFBS | <0.552 | <0.608 | <0.613
PFHxS | <0.721 | <0.794 | <0.801
PFOS | 6.33 | 6.26 | 6.11 |  |  |  | 5.9 | ± | 0.39
PFOSA | <0.848 | <0.934 | <0.942
of